## 200623 - 100g Standard Mass
Mass Balance Test
Device Under Test: | OS_Nano_Balance (3D printed balance), Using TAL221 100g load cell
Lab Balance: | N/A Using standard masses
DUT Calibration: | Calibrated with 100 g standard mass
| 13914.92 | g/div
PuTTY
[8 empty] | -0.00462 | g
[8 empty] | -0.00362 | g
|  |  | Nominal | Average | Std | Count | Addresses | -0.00139 | g
|  |  | 1 | 1.0002 | 0.0025389 | 28 | R22:R49 | -0.00169 | g
|  |  | 2 | 1.9941 | 0.0021696 | 29 | R64:R92 | 0.00165 | g
|  |  | 5 | 4.9864 | 0.0017971 | 30 | R102:R131 | 0.00067 | g
|  |  | 10 | 9.9908 | 0.0017282 | 32 | R140:R171 | 0.00118 | g
|  |  | 20 | 19.997 | 0.0026241 | 29 | R182:R210 | -6e-05 | g
|  |  | 50 | 50.0405 | 0.0056585 | 31 | R221:R251 | 0.00119 | g
|  |  | 1 | 1.0194 | 0.0016537 | 34 | R320:R353 | 0.00249 | g
|  |  | 2 | 2.0146 | 0.0012197 | 27 | R369:R395 | 0.94735 | g
|  |  | 5 | 5.0065 | 0.0012887 | 29 | R405:R433 | 1.0025 | g
|  |  | 10 | 10.0083 | 0.0019213 | 36 | R444:R479 | 1.0055 | g
|  |  | 20 | 20.0196 | 0.0030292 | 35 | R490:R524 | 1.002 | g
|  |  | 50 | 50.0556 | 0.0059258 | 32 | R534:R565 | 1.0039 | g
[8 empty] | 1.0024 | g
[8 empty] | 1.0019 | g
[8 empty] | 1.002 | g
|  |  |  | Avg | 0.0026296 |  | 0.00263% | 0.99801 | g
|  |  |  | Max | 0.0059258 |  |  | 0.99988 | g
[8 empty] | 0.99769 | g
|  |  | 100 | 100.0767 | 0.010993 | 34 | R267:R300 | 0.99878 | g
|  |  | 100 | 100.0869 | 0.01177 | 28 | R590:R617 | 0.99613 | g
[8 empty] | 0.99817 | g
[8 empty] | 1.0019 | g
[8 empty] | 1.0012 | g
[8 empty] | 1.003 | g
[8 empty] | 1.0036 | g
[8 empty] | 1.0021 | g
[8 empty] | 1.0013 | g
[8 empty] | 0.99666 | g
[8 empty] | 0.99554 | g
[8 empty] | 0.99851 | g
[8 empty] | 0.99661 | g
[8 empty] | 1.0008 | g
[8 empty] | 0.99942 | g
[8 empty] | 0.99858 | g
[8 empty] | 0.99845 | g
[8 empty] | 0.99947 | g
[8 empty] | 1.1359 | g
[8 empty] | 0.01222 | g
[8 empty] | -0.00237 | g
[8 empty] | -0.00142 | g
[8 empty] | -0.00316 | g
[8 empty] | -0.00131 | g
[8 empty] | -0.00075 | g
[8 empty] | -0.00137 | g
[8 empty] | -0.00451 | g
[8 empty] | -0.00584 | g
[8 empty] | -0.00695 | g
[8 empty] | -0.00574 | g
[8 empty] | -0.00663 | g
[8 empty] | 0.82138 | g
[8 empty] | 1.992 | g
[8 empty] | 1.9884 | g
[8 empty] | 1.9911 | g
[8 empty] | 1.9902 | g
[8 empty] | 1.9911 | g
[8 empty] | 1.9916 | g
[8 empty] | 1.9946 | g
[8 empty] | 1.9949 | g
[8 empty] | 1.9942 | g
[8 empty] | 1.9936 | g
[8 empty] | 1.9952 | g
[8 empty] | 1.9953 | g
[8 empty] | 1.9954 | g
[8 empty] | 1.9964 | g
[8 empty] | 1.995 | g
[8 empty] | 1.9953 | g
[8 empty] | 1.9956 | g
[8 empty] | 1.9956 | g
[8 empty] | 1.995 | g
[8 empty] | 1.9942 | g
[8 empty] | 1.9934 | g
[8 empty] | 1.9935 | g
[8 empty] | 1.9935 | g
[8 empty] | 1.9942 | g
[8 empty] | 1.9957 | g
[8 empty] | 1.9928 | g
[8 empty] | 1.9961 | g
[8 empty] | 1.9974 | g
[8 empty] | 1.9988 | g
[8 empty] | 1.5908 | g
[8 empty] | -0.0078 | g
[8 empty] | -0.00888 | g
[8 empty] | -0.00739 | g
[8 empty] | -0.00714 | g
[8 empty] | -0.00618 | g
[8 empty] | -0.00536 | g
[8 empty] | 3.0689 | g
[8 empty] | 4.9821 | g
[8 empty] | 4.9842 | g
[8 empty] | 4.9825 | g
[8 empty] | 4.9847 | g
[8 empty] | 4.9828 | g
[8 empty] | 4.9841 | g
[8 empty] | 4.9859 | g
[8 empty] | 4.9863 | g
[8 empty] | 4.9866 | g
[8 empty] | 4.9858 | g
[8 empty] | 4.9857 | g
[8 empty] | 4.985 | g
[8 empty] | 4.9846 | g
[8 empty] | 4.9871 | g
[8 empty] | 4.9872 | g
[8 empty] | 4.9867 | g
[8 empty] | 4.9849 | g
[8 empty] | 4.9897 | g
[8 empty] | 4.9871 | g
[8 empty] | 4.9881 | g
[8 empty] | 4.9887 | g
[8 empty] | 4.9882 | g
[8 empty] | 4.9874 | g
[8 empty] | 4.9852 | g
[8 empty] | 4.9863 | g
[8 empty] | 4.9868 | g
[8 empty] | 4.9876 | g
[8 empty] | 4.9878 | g
[8 empty] | 4.9891 | g
[8 empty] | 4.9882 | g
[8 empty] | 4.9889 | g
[8 empty] | 3.5467 | g
[8 empty] | -0.00096 | g
[8 empty] | -0.00776 | g
[8 empty] | -0.0082 | g
[8 empty] | -0.00803 | g
[8 empty] | -0.00802 | g
[8 empty] | 0.12763 | g
[8 empty] | 9.9859 | g
[8 empty] | 9.9867 | g
[8 empty] | 9.9878 | g
[8 empty] | 9.9882 | g
[8 empty] | 9.9889 | g
[8 empty] | 9.9898 | g
[8 empty] | 9.9892 | g
[8 empty] | 9.9897 | g
[8 empty] | 9.9899 | g
[8 empty] | 9.99 | g
[8 empty] | 9.99 | g
[8 empty] | 9.9907 | g
[8 empty] | 9.9898 | g
[8 empty] | 9.9919 | g
[8 empty] | 9.9906 | g
[8 empty] | 9.9904 | g
[8 empty] | 9.991 | g
[8 empty] | 9.9929 | g
[8 empty] | 9.9922 | g
[8 empty] | 9.9921 | g
[8 empty] | 9.9921 | g
[8 empty] | 9.9914 | g
[8 empty] | 9.9917 | g
[8 empty] | 9.9905 | g
[8 empty] | 9.99 | g
[8 empty] | 9.9913 | g
[8 empty] | 9.9924 | g
[8 empty] | 9.9915 | g
[8 empty] | 9.9923 | g
[8 empty] | 9.9898 | g
[8 empty] | 9.9941 | g
[8 empty] | 9.993 | g
[8 empty] | 9.9948 | g
[8 empty] | 10.2865 | g
[8 empty] | 0.36868 | g
[8 empty] | -0.00103 | g
[8 empty] | -0.00086 | g
[8 empty] | -0.00185 | g
[8 empty] | -0.00421 | g
[8 empty] | -0.00407 | g
[8 empty] | -0.00232 | g
[8 empty] | -0.00382 | g
[8 empty] | 13.6017 | g
[8 empty] | 19.9907 | g
[8 empty] | 19.9905 | g
[8 empty] | 19.9924 | g
[8 empty] | 19.9927 | g
[8 empty] | 19.9956 | g
[8 empty] | 19.9958 | g
[8 empty] | 19.9961 | g
[8 empty] | 19.9973 | g
[8 empty] | 19.9953 | g
[8 empty] | 19.9968 | g
[8 empty] | 19.9982 | g
[8 empty] | 19.9972 | g
[8 empty] | 19.9967 | g
[8 empty] | 19.9999 | g
[8 empty] | 19.9992 | g
[8 empty] | 20.0001 | g
[8 empty] | 19.9988 | g
[8 empty] | 20.0009 | g
[8 empty] | 19.9964 | g
[8 empty] | 19.9978 | g
[8 empty] | 19.9987 | g
[8 empty] | 19.9972 | g
[8 empty] | 19.9988 | g
[8 empty] | 19.9982 | g
[8 empty] | 19.9978 | g
[8 empty] | 19.9979 | g
[8 empty] | 19.9981 | g
[8 empty] | 19.9985 | g
[8 empty] | 20.0006 | g
[8 empty] | 20.0872 | g
[8 empty] | 4.4423 | g
[8 empty] | -0.00126 | g
[8 empty] | -0.0002 | g
[8 empty] | -0.00019 | g
[8 empty] | -0.00041 | g
[8 empty] | -0.00113 | g
[8 empty] | -0.00064 | g
[8 empty] | 16.5119 | g
[8 empty] | 49.9196 | g
[8 empty] | 50.0256 | g
[8 empty] | 50.0289 | g
[8 empty] | 50.0303 | g
[8 empty] | 50.0338 | g
[8 empty] | 50.0357 | g
[8 empty] | 50.0348 | g
[8 empty] | 50.0365 | g
[8 empty] | 50.0377 | g
[8 empty] | 50.0381 | g
[8 empty] | 50.0375 | g
[8 empty] | 50.0384 | g
[8 empty] | 50.0405 | g
[8 empty] | 50.0385 | g
[8 empty] | 50.0415 | g
[8 empty] | 50.042 | g
[8 empty] | 50.0405 | g
[8 empty] | 50.0417 | g
[8 empty] | 50.041 | g
[8 empty] | 50.0419 | g
[8 empty] | 50.0428 | g
[8 empty] | 50.0443 | g
[8 empty] | 50.0445 | g
[8 empty] | 50.0456 | g
[8 empty] | 50.0468 | g
[8 empty] | 50.0456 | g
[8 empty] | 50.0452 | g
[8 empty] | 50.0465 | g
[8 empty] | 50.0457 | g
[8 empty] | 50.0478 | g
[8 empty] | 50.0469 | g
[8 empty] | 50.0489 | g
[8 empty] | 45.5815 | g
[8 empty] | 0.03114 | g
[8 empty] | 0.02226 | g
[8 empty] | 0.02045 | g
[8 empty] | 0.01667 | g
[8 empty] | 0.01572 | g
[8 empty] | 0.01591 | g
[8 empty] | 0.01341 | g
[8 empty] | 0.01236 | g
[8 empty] | 0.01248 | g
[8 empty] | 0.00974 | g
[8 empty] | 0.00903 | g
[8 empty] | 0.00994 | g
[8 empty] | 5.1887 | g
[8 empty] | 96.9146 | g
[8 empty] | 100.0434 | g
[8 empty] | 100.0518 | g
[8 empty] | 100.0571 | g
[8 empty] | 100.0606 | g
[8 empty] | 100.0639 | g
[8 empty] | 100.0676 | g
[8 empty] | 100.0696 | g
[8 empty] | 100.0694 | g
[8 empty] | 100.0719 | g
[8 empty] | 100.0724 | g
[8 empty] | 100.0759 | g
[8 empty] | 100.0749 | g
[8 empty] | 100.0748 | g
[8 empty] | 100.0773 | g
[8 empty] | 100.0759 | g
[8 empty] | 100.0777 | g
[8 empty] | 100.0782 | g
[8 empty] | 100.0791 | g
[8 empty] | 100.0797 | g
[8 empty] | 100.0806 | g
[8 empty] | 100.0837 | g
[8 empty] | 100.0818 | g
[8 empty] | 100.0825 | g
[8 empty] | 100.0841 | g
[8 empty] | 100.0836 | g
[8 empty] | 100.0856 | g
[8 empty] | 100.0873 | g
[8 empty] | 100.0855 | g
[8 empty] | 100.0884 | g
[8 empty] | 100.0863 | g
[8 empty] | 100.0888 | g
[8 empty] | 100.0878 | g
[8 empty] | 100.0891 | g
[8 empty] | 100.09 | g
[8 empty] | 101.7625 | g
[8 empty] | 18.7211 | g
[8 empty] | 0.05402 | g
[8 empty] | 0.04546 | g
[8 empty] | 0.0406 | g
[8 empty] | 0.03839 | g
[8 empty] | 0.0337 | g
[8 empty] | 0.03291 | g
[8 empty] | 0.03345 | g
[8 empty] | 0.0288 | g
[8 empty] | 0.02723 | g
[8 empty] | 0.02864 | g
[8 empty] | 0.02646 | g
[8 empty] | 0.02573 | g
[8 empty] | 0.02306 | g
[8 empty] | 0.02352 | g
[8 empty] | 0.02325 | g
[8 empty] | 0.02303 | g
[8 empty] | 0.64562 | g
[8 empty] | 1.0224 | g
[8 empty] | 1.0223 | g
[8 empty] | 1.0221 | g
[8 empty] | 1.0203 | g
[8 empty] | 1.0214 | g
[8 empty] | 1.0229 | g
[8 empty] | 1.0205 | g
[8 empty] | 1.0187 | g
[8 empty] | 1.0207 | g
[8 empty] | 1.0203 | g
[8 empty] | 1.0198 | g
[8 empty] | 1.0218 | g
[8 empty] | 1.0195 | g
[8 empty] | 1.0195 | g
[8 empty] | 1.0193 | g
[8 empty] | 1.0192 | g
[8 empty] | 1.0178 | g
[8 empty] | 1.0186 | g
[8 empty] | 1.0181 | g
[8 empty] | 1.0194 | g
[8 empty] | 1.0177 | g
[8 empty] | 1.0183 | g
[8 empty] | 1.019 | g
[8 empty] | 1.0163 | g
[8 empty] | 1.02 | g
[8 empty] | 1.0181 | g
[8 empty] | 1.0179 | g
[8 empty] | 1.0181 | g
[8 empty] | 1.0189 | g
[8 empty] | 1.0173 | g
[8 empty] | 1.0168 | g
[8 empty] | 1.0201 | g
[8 empty] | 1.0186 | g
[8 empty] | 1.0174 | g
[8 empty] | 0.88315 | g
[8 empty] | 0.01421 | g
[8 empty] | 0.01559 | g
[8 empty] | 0.01371 | g
[8 empty] | 0.01347 | g
[8 empty] | 0.01393 | g
[8 empty] | 0.0127 | g
[8 empty] | 0.01289 | g
[8 empty] | 0.01414 | g
[8 empty] | 0.01161 | g
[8 empty] | 0.01535 | g
[8 empty] | 0.01105 | g
[8 empty] | 0.01104 | g
[8 empty] | 1.2355 | g
[8 empty] | 2.0124 | g
[8 empty] | 2.0144 | g
[8 empty] | 2.0143 | g
[8 empty] | 2.0161 | g
[8 empty] | 2.0145 | g
[8 empty] | 2.0153 | g
[8 empty] | 2.0155 | g
[8 empty] | 2.0155 | g
[8 empty] | 2.0137 | g
[8 empty] | 2.0148 | g
[8 empty] | 2.0155 | g
[8 empty] | 2.012 | g
[8 empty] | 2.012 | g
[8 empty] | 2.0131 | g
[8 empty] | 2.0131 | g
[8 empty] | 2.0142 | g
[8 empty] | 2.0149 | g
[8 empty] | 2.0156 | g
[8 empty] | 2.0136 | g
[8 empty] | 2.0124 | g
[8 empty] | 2.0143 | g
[8 empty] | 2.0167 | g
[8 empty] | 2.0154 | g
[8 empty] | 2.0162 | g
[8 empty] | 2.0148 | g
[8 empty] | 2.0145 | g
[8 empty] | 2.0148 | g
[8 empty] | 2.0157 | g
[8 empty] | 1.9944 | g
[8 empty] | 0.01529 | g
[8 empty] | 0.01334 | g
[8 empty] | 0.0101 | g
[8 empty] | 0.01145 | g
[8 empty] | 0.01191 | g
[8 empty] | 0.01066 | g
[8 empty] | 0.0121 | g
[8 empty] | 1.572 | g
[8 empty] | 5.0059 | g
[8 empty] | 5.0053 | g
[8 empty] | 5.0056 | g
[8 empty] | 5.006 | g
[8 empty] | 5.0067 | g
[8 empty] | 5.004 | g
[8 empty] | 5.0054 | g
[8 empty] | 5.0051 | g
[8 empty] | 5.0051 | g
[8 empty] | 5.0063 | g
[8 empty] | 5.0056 | g
[8 empty] | 5.0073 | g
[8 empty] | 5.0049 | g
[8 empty] | 5.0075 | g
[8 empty] | 5.008 | g
[8 empty] | 5.0065 | g
[8 empty] | 5.0067 | g
[8 empty] | 5.0074 | g
[8 empty] | 5.0064 | g
[8 empty] | 5.0058 | g
[8 empty] | 5.0065 | g
[8 empty] | 5.0046 | g
[8 empty] | 5.0084 | g
[8 empty] | 5.0088 | g
[8 empty] | 5.0091 | g
[8 empty] | 5.0085 | g
[8 empty] | 5.0077 | g
[8 empty] | 5.0062 | g
[8 empty] | 5.0059 | g
[8 empty] | 3.7344 | g
[8 empty] | 0.01705 | g
[8 empty] | 0.01207 | g
[8 empty] | 0.01192 | g
[8 empty] | 0.01079 | g
[8 empty] | 0.00965 | g
[8 empty] | 0.01015 | g
[8 empty] | 0.00985 | g
[8 empty] | 0.00911 | g
[8 empty] | 6.9425 | g
[8 empty] | 10.0032 | g
[8 empty] | 10.0052 | g
[8 empty] | 10.0065 | g
[8 empty] | 10.0075 | g
[8 empty] | 10.0083 | g
[8 empty] | 10.0049 | g
[8 empty] | 10.0065 | g
[8 empty] | 10.0077 | g
[8 empty] | 10.0084 | g
[8 empty] | 10.0087 | g
[8 empty] | 10.0063 | g
[8 empty] | 10.008 | g
[8 empty] | 10.0078 | g
[8 empty] | 10.0078 | g
[8 empty] | 10.0073 | g
[8 empty] | 10.0069 | g
[8 empty] | 10.0074 | g
[8 empty] | 10.0104 | g
[8 empty] | 10.0073 | g
[8 empty] | 10.0088 | g
[8 empty] | 10.0057 | g
[8 empty] | 10.0086 | g
[8 empty] | 10.0076 | g
[8 empty] | 10.0104 | g
[8 empty] | 10.0108 | g
[8 empty] | 10.0096 | g
[8 empty] | 10.0114 | g
[8 empty] | 10.0104 | g
[8 empty] | 10.0085 | g
[8 empty] | 10.0101 | g
[8 empty] | 10.01 | g
[8 empty] | 10.0119 | g
[8 empty] | 10.0083 | g
[8 empty] | 10.0096 | g
[8 empty] | 10.0112 | g
[8 empty] | 10.0082 | g
[8 empty] | 11.1708 | g
[8 empty] | 11.2566 | g
[8 empty] | 11.1605 | g
[8 empty] | 0.01713 | g
[8 empty] | 0.01495 | g
[8 empty] | 0.0135 | g
[8 empty] | 0.01436 | g
[8 empty] | 0.01355 | g
[8 empty] | 0.01252 | g
[8 empty] | 8.8165 | g
[8 empty] | 20.0125 | g
[8 empty] | 20.0144 | g
[8 empty] | 20.0138 | g
[8 empty] | 20.0157 | g
[8 empty] | 20.016 | g
[8 empty] | 20.0173 | g
[8 empty] | 20.0152 | g
[8 empty] | 20.018 | g
[8 empty] | 20.0172 | g
[8 empty] | 20.0212 | g
[8 empty] | 20.0191 | g
[8 empty] | 20.0182 | g
[8 empty] | 20.0181 | g
[8 empty] | 20.0187 | g
[8 empty] | 20.0187 | g
[8 empty] | 20.0203 | g
[8 empty] | 20.0206 | g
[8 empty] | 20.0209 | g
[8 empty] | 20.02 | g
[8 empty] | 20.0183 | g
[8 empty] | 20.0196 | g
[8 empty] | 20.0216 | g
[8 empty] | 20.0206 | g
[8 empty] | 20.022 | g
[8 empty] | 20.0206 | g
[8 empty] | 20.0221 | g
[8 empty] | 20.02 | g
[8 empty] | 20.0234 | g
[8 empty] | 20.0225 | g
[8 empty] | 20.0236 | g
[8 empty] | 20.0233 | g
[8 empty] | 20.0214 | g
[8 empty] | 20.0251 | g
[8 empty] | 20.0229 | g
[8 empty] | 20.0241 | g
[8 empty] | 20.7026 | g
[8 empty] | 4.169 | g
[8 empty] | 0.02236 | g
[8 empty] | 0.02293 | g
[8 empty] | 0.02084 | g
[8 empty] | 0.02031 | g
[8 empty] | 0.01853 | g
[8 empty] | 0.01764 | g
[8 empty] | 26.8555 | g
[8 empty] | 50.0369 | g
[8 empty] | 50.0443 | g
[8 empty] | 50.0469 | g
[8 empty] | 50.0481 | g
[8 empty] | 50.047 | g
[8 empty] | 50.0498 | g
[8 empty] | 50.0496 | g
[8 empty] | 50.0541 | g
[8 empty] | 50.0547 | g
[8 empty] | 50.0545 | g
[8 empty] | 50.0558 | g
[8 empty] | 50.0546 | g
[8 empty] | 50.0564 | g
[8 empty] | 50.0558 | g
[8 empty] | 50.0566 | g
[8 empty] | 50.055 | g
[8 empty] | 50.0562 | g
[8 empty] | 50.0559 | g
[8 empty] | 50.0575 | g
[8 empty] | 50.0589 | g
[8 empty] | 50.0586 | g
[8 empty] | 50.0594 | g
[8 empty] | 50.0593 | g
[8 empty] | 50.0624 | g
[8 empty] | 50.059 | g
[8 empty] | 50.0619 | g
[8 empty] | 50.0608 | g
[8 empty] | 50.0603 | g
[8 empty] | 50.0602 | g
[8 empty] | 50.0616 | g
[8 empty] | 50.0639 | g
[8 empty] | 50.0619 | g
[8 empty] | 50.3149 | g
[8 empty] | 11.474 | g
[8 empty] | 0.03835 | g
[8 empty] | 0.03451 | g
[8 empty] | 0.03202 | g
[8 empty] | 0.02917 | g
[8 empty] | 0.02722 | g
[8 empty] | 0.02793 | g
[8 empty] | 0.02527 | g
[8 empty] | 0.02563 | g
[8 empty] | 0.02386 | g
[8 empty] | 0.02216 | g
[8 empty] | 0.02163 | g
[8 empty] | 0.02235 | g
[8 empty] | 0.0221 | g
[8 empty] | 0.0216 | g
[8 empty] | 0.01929 | g
[8 empty] | 0.02163 | g
[8 empty] | 0.02094 | g
[8 empty] | 0.01925 | g
[8 empty] | 0.02137 | g
[8 empty] | 0.01931 | g
[8 empty] | 22.3823 | g
[8 empty] | 100.0045 | g
[8 empty] | 100.0569 | g
[8 empty] | 100.0641 | g
[8 empty] | 100.0688 | g
[8 empty] | 100.0717 | g
[8 empty] | 100.0749 | g
[8 empty] | 100.0749 | g
[8 empty] | 100.0776 | g
[8 empty] | 100.0795 | g
[8 empty] | 100.0813 | g
[8 empty] | 100.082 | g
[8 empty] | 100.085 | g
[8 empty] | 100.0854 | g
[8 empty] | 100.0886 | g
[8 empty] | 100.0901 | g
[8 empty] | 100.0923 | g
[8 empty] | 100.0955 | g
[8 empty] | 100.0928 | g
[8 empty] | 100.0959 | g
[8 empty] | 100.0956 | g
[8 empty] | 100.0957 | g
[8 empty] | 100.0959 | g
[8 empty] | 100.0968 | g
[8 empty] | 100.0976 | g
[8 empty] | 100.0992 | g
[8 empty] | 100.0993 | g
[8 empty] | 100.0991 | g
[8 empty] | 100.0985 | g
[8 empty] | 100.0987 | g
[8 empty] | 100.1021 | g
[8 empty] | 99.5626 | g
[8 empty] | 4.1985 | g
[8 empty] | 0.06292 | g
[8 empty] | 0.05772 | g
[8 empty] | 0.05038 | g
[8 empty] | 0.04919 | g
[8 empty] | 0.04635 | g
[8 empty] | 0.04177 | g
[8 empty] | 0.04193 | g
[8 empty] | 0.04055 | g
[8 empty] | 0.03979 | g
[8 empty] | 0.03647 | g
[8 empty] | 0.03679 | g
[8 empty] | 0.03474 | g

## 200623 - 500g Standard Mass
Mass Balance Test
Device Under Test: | OS_Nano_Balance (3D printed balance), Using TAL221 500g load cell
Lab Balance: | N/A Using standard masses
DUT Calibration: | Calibrated with 100 g standard mass
| 3243.81 | g/div
PuTTY
[8 empty] | -0.013 | g
[8 empty] | 0.009 | g
|  |  | Nominal | Average | Std | Count | Addresses | -0.004 | g
|  |  | 1 | 0.99056 | 0.0051014 | 27 | R24:R50 | -0.004 | g
|  |  | 2 | 1.9954 | 0.0042091 | 23 | R63:R85 | 0.001 | g
|  |  | 5 | 4.982 | 0.0061669 | 32 | R97:R128 | 0.007 | g
|  |  | 10 | 9.9827 | 0.017558 | 29 | R142:R170 | 0.004 | g
|  |  | 20 | 19.9888 | 0.005311 | 26 | R181:R206 | -0.007 | g
|  |  | 50 | 49.9845 | 0.005641 | 28 | R217:R244 | -0.004 | g
|  |  | 100 | 99.9479 | 0.0064877 | 30 | R262:R291 | -0.008 | g
|  |  | 1 | 0.99467 | 0.0063421 | 27 | R319:R345 | -0.011 | g
|  |  | 2 | 1.9933 | 0.0042475 | 25 | R358:R382 | -0.01 | g
|  |  | 5 | 4.9868 | 0.0047838 | 31 | R397:R427 | 0.557 | g
|  |  | 10 | 9.9909 | 0.0049391 | 27 | R442:R468 | 0.986 | g
|  |  | 20 | 19.9893 | 0.0052607 | 26 | R480:R505 | 0.986 | g
|  |  | 50 | 49.9922 | 0.0072095 | 26 | R515:R540 | 0.996 | g
|  |  | 100 | 99.9683 | 0.015469 | 26 | R556:R581 | 0.988 | g
[8 empty] | 0.993 | g
|  |  |  | Avg | 0.0070519 |  | 0.00141% | 0.99 | g
|  |  |  | Max | 0.017558 |  |  | 0.992 | g
[8 empty] | 0.995 | g
[8 empty] | 0.992 | g
[8 empty] | 0.989 | g
[8 empty] | 0.993 | g
[8 empty] | 1.002 | g
[8 empty] | 0.999 | g
[8 empty] | 0.994 | g
[8 empty] | 0.993 | g
[8 empty] | 0.994 | g
[8 empty] | 0.993 | g
[8 empty] | 0.99 | g
[8 empty] | 0.977 | g
[8 empty] | 0.986 | g
[8 empty] | 0.986 | g
[8 empty] | 0.989 | g
[8 empty] | 0.99 | g
[8 empty] | 0.984 | g
[8 empty] | 0.995 | g
[8 empty] | 0.983 | g
[8 empty] | 0.99 | g
[8 empty] | 0.836 | g
[8 empty] | -0.021 | g
[8 empty] | -0.018 | g
[8 empty] | -0.011 | g
[8 empty] | -0.006 | g
[8 empty] | -0.016 | g
[8 empty] | -0.008 | g
[8 empty] | -0.017 | g
[8 empty] | -0.016 | g
[8 empty] | -0.013 | g
[8 empty] | -0.017 | g
[8 empty] | 1.114 | g
[8 empty] | 1.998 | g
[8 empty] | 1.992 | g
[8 empty] | 1.993 | g
[8 empty] | 2.003 | g
[8 empty] | 1.994 | g
[8 empty] | 1.997 | g
[8 empty] | 2.004 | g
[8 empty] | 1.996 | g
[8 empty] | 1.988 | g
[8 empty] | 1.994 | g
[8 empty] | 2 | g
[8 empty] | 1.997 | g
[8 empty] | 1.997 | g
[8 empty] | 1.991 | g
[8 empty] | 1.992 | g
[8 empty] | 2.001 | g
[8 empty] | 2 | g
[8 empty] | 1.995 | g
[8 empty] | 1.991 | g
[8 empty] | 1.996 | g
[8 empty] | 1.991 | g
[8 empty] | 1.995 | g
[8 empty] | 1.989 | g
[8 empty] | 2.698 | g
[8 empty] | 0.153 | g
[8 empty] | -0.018 | g
[8 empty] | -0.009 | g
[8 empty] | -0.012 | g
[8 empty] | -0.011 | g
[8 empty] | -0.007 | g
[8 empty] | -0.009 | g
[8 empty] | -0.004 | g
[8 empty] | -0.011 | g
[8 empty] | 1.821 | g
[8 empty] | 4.982 | g
[8 empty] | 4.981 | g
[8 empty] | 4.994 | g
[8 empty] | 4.991 | g
[8 empty] | 4.988 | g
[8 empty] | 4.991 | g
[8 empty] | 4.997 | g
[8 empty] | 4.989 | g
[8 empty] | 4.986 | g
[8 empty] | 4.981 | g
[8 empty] | 4.989 | g
[8 empty] | 4.982 | g
[8 empty] | 4.987 | g
[8 empty] | 4.979 | g
[8 empty] | 4.974 | g
[8 empty] | 4.977 | g
[8 empty] | 4.975 | g
[8 empty] | 4.98 | g
[8 empty] | 4.977 | g
[8 empty] | 4.977 | g
[8 empty] | 4.977 | g
[8 empty] | 4.983 | g
[8 empty] | 4.982 | g
[8 empty] | 4.977 | g
[8 empty] | 4.982 | g
[8 empty] | 4.985 | g
[8 empty] | 4.975 | g
[8 empty] | 4.983 | g
[8 empty] | 4.972 | g
[8 empty] | 4.973 | g
[8 empty] | 4.978 | g
[8 empty] | 4.981 | g
[8 empty] | 2.422 | g
[8 empty] | -0.019 | g
[8 empty] | -0.016 | g
[8 empty] | -0.021 | g
[8 empty] | -0.028 | g
[8 empty] | -0.015 | g
[8 empty] | -0.015 | g
[8 empty] | -0.011 | g
[8 empty] | -0.019 | g
[8 empty] | -0.017 | g
[8 empty] | -0.018 | g
[8 empty] | -0.013 | g
[8 empty] | 6.075 | g
[8 empty] | 9.983 | g
[8 empty] | 9.983 | g
[8 empty] | 9.989 | g
[8 empty] | 9.987 | g
[8 empty] | 9.98 | g
[8 empty] | 9.978 | g
[8 empty] | 9.981 | g
[8 empty] | 9.982 | g
[8 empty] | 9.969 | g
[8 empty] | 9.983 | g
[8 empty] | 9.979 | g
[8 empty] | 9.971 | g
[8 empty] | 9.98 | g
[8 empty] | 9.982 | g
[8 empty] | 9.989 | g
[8 empty] | 9.992 | g
[8 empty] | 9.996 | g
[8 empty] | 9.989 | g
[8 empty] | 9.985 | g
[8 empty] | 9.994 | g
[8 empty] | 9.994 | g
[8 empty] | 9.989 | g
[8 empty] | 9.99 | g
[8 empty] | 9.999 | g
[8 empty] | 9.985 | g
[8 empty] | 9.99 | g
[8 empty] | 9.991 | g
[8 empty] | 9.991 | g
[8 empty] | 9.897 | g
[8 empty] | 7.101 | g
[8 empty] | -0.002 | g
[8 empty] | -0.005 | g
[8 empty] | -0.004 | g
[8 empty] | 0.001 | g
[8 empty] | 0.005 | g
[8 empty] | -0.002 | g
[8 empty] | 0.011 | g
[8 empty] | 0.004 | g
[8 empty] | 11.882 | g
[8 empty] | 19.996 | g
[8 empty] | 19.988 | g
[8 empty] | 19.991 | g
[8 empty] | 19.996 | g
[8 empty] | 19.995 | g
[8 empty] | 19.993 | g
[8 empty] | 19.995 | g
[8 empty] | 19.994 | g
[8 empty] | 19.994 | g
[8 empty] | 19.993 | g
[8 empty] | 19.995 | g
[8 empty] | 19.989 | g
[8 empty] | 19.987 | g
[8 empty] | 19.986 | g
[8 empty] | 19.986 | g
[8 empty] | 19.991 | g
[8 empty] | 19.985 | g
[8 empty] | 19.989 | g
[8 empty] | 19.988 | g
[8 empty] | 19.984 | g
[8 empty] | 19.984 | g
[8 empty] | 19.98 | g
[8 empty] | 19.985 | g
[8 empty] | 19.976 | g
[8 empty] | 19.98 | g
[8 empty] | 19.99 | g
[8 empty] | 20.738 | g
[8 empty] | 1.909 | g
[8 empty] | -0.008 | g
[8 empty] | -0.004 | g
[8 empty] | -0.012 | g
[8 empty] | 0.001 | g
[8 empty] | -0.008 | g
[8 empty] | -0.013 | g
[8 empty] | -0.004 | g
[8 empty] | 40.678 | g
[8 empty] | 49.977 | g
[8 empty] | 49.99 | g
[8 empty] | 49.984 | g
[8 empty] | 49.992 | g
[8 empty] | 49.979 | g
[8 empty] | 49.989 | g
[8 empty] | 49.981 | g
[8 empty] | 49.981 | g
[8 empty] | 49.975 | g
[8 empty] | 49.983 | g
[8 empty] | 49.982 | g
[8 empty] | 49.992 | g
[8 empty] | 49.994 | g
[8 empty] | 49.985 | g
[8 empty] | 49.982 | g
[8 empty] | 49.984 | g
[8 empty] | 49.984 | g
[8 empty] | 49.985 | g
[8 empty] | 49.984 | g
[8 empty] | 49.975 | g
[8 empty] | 49.982 | g
[8 empty] | 49.982 | g
[8 empty] | 49.985 | g
[8 empty] | 49.991 | g
[8 empty] | 49.991 | g
[8 empty] | 49.98 | g
[8 empty] | 49.979 | g
[8 empty] | 49.998 | g
[8 empty] | 19.469 | g
[8 empty] | -0.006 | g
[8 empty] | -0.004 | g
[8 empty] | -0.002 | g
[8 empty] | -0.011 | g
[8 empty] | -0.003 | g
[8 empty] | -0.014 | g
[8 empty] | -0.014 | g
[8 empty] | -0.01 | g
[8 empty] | -0.01 | g
[8 empty] | -0.015 | g
[8 empty] | -0.017 | g
[8 empty] | -0.014 | g
[8 empty] | -0.01 | g
[8 empty] | -0.012 | g
[8 empty] | 6.297 | g
[8 empty] | 98.719 | g
[8 empty] | 99.936 | g
[8 empty] | 99.946 | g
[8 empty] | 99.945 | g
[8 empty] | 99.944 | g
[8 empty] | 99.937 | g
[8 empty] | 99.943 | g
[8 empty] | 99.94 | g
[8 empty] | 99.933 | g
[8 empty] | 99.949 | g
[8 empty] | 99.944 | g
[8 empty] | 99.951 | g
[8 empty] | 99.952 | g
[8 empty] | 99.951 | g
[8 empty] | 99.964 | g
[8 empty] | 99.954 | g
[8 empty] | 99.946 | g
[8 empty] | 99.948 | g
[8 empty] | 99.95 | g
[8 empty] | 99.948 | g
[8 empty] | 99.95 | g
[8 empty] | 99.949 | g
[8 empty] | 99.954 | g
[8 empty] | 99.943 | g
[8 empty] | 99.953 | g
[8 empty] | 99.953 | g
[8 empty] | 99.957 | g
[8 empty] | 99.948 | g
[8 empty] | 99.957 | g
[8 empty] | 99.943 | g
[8 empty] | 99.949 | g
[8 empty] | 99.402 | g
[8 empty] | 71.465 | g
[8 empty] | 0.011 | g
[8 empty] | -0.001 | g
[8 empty] | -0.008 | g
[8 empty] | 0.002 | g
[8 empty] | -0.001 | g
[8 empty] | 0.004 | g
[8 empty] | -0.002 | g
[8 empty] | 0.004 | g
[8 empty] | 0.003 | g
[8 empty] | -0.008 | g
[8 empty] | -0.003 | g
[8 empty] | -0.013 | g
[8 empty] | -0.006 | g
[8 empty] | -0.005 | g
[8 empty] | -0.003 | g
[8 empty] | -0.009 | g
[8 empty] | -0.006 | g
[8 empty] | -0.006 | g
[8 empty] | 0.002 | g
[8 empty] | -0.009 | g
[8 empty] | -0.002 | g
[8 empty] | -0.011 | g
[8 empty] | -0.012 | g
[8 empty] | 0 | g
[8 empty] | 0.016 | g
[8 empty] | 0.997 | g
[8 empty] | 1.005 | g
[8 empty] | 1.005 | g
[8 empty] | 0.999 | g
[8 empty] | 1.007 | g
[8 empty] | 0.994 | g
[8 empty] | 0.999 | g
[8 empty] | 0.992 | g
[8 empty] | 0.992 | g
[8 empty] | 0.993 | g
[8 empty] | 0.993 | g
[8 empty] | 0.99 | g
[8 empty] | 0.993 | g
[8 empty] | 1.004 | g
[8 empty] | 0.992 | g
[8 empty] | 0.99 | g
[8 empty] | 1.006 | g
[8 empty] | 0.991 | g
[8 empty] | 0.985 | g
[8 empty] | 0.982 | g
[8 empty] | 0.986 | g
[8 empty] | 0.99 | g
[8 empty] | 0.995 | g
[8 empty] | 0.99 | g
[8 empty] | 0.995 | g
[8 empty] | 0.995 | g
[8 empty] | 0.996 | g
[8 empty] | 0.376 | g
[8 empty] | -0.008 | g
[8 empty] | -0.01 | g
[8 empty] | -0.002 | g
[8 empty] | -0.005 | g
[8 empty] | -0.008 | g
[8 empty] | -0.021 | g
[8 empty] | -0.008 | g
[8 empty] | -0.002 | g
[8 empty] | -0.01 | g
[8 empty] | -0.011 | g
[8 empty] | 3.551 | g
[8 empty] | 1.993 | g
[8 empty] | 1.99 | g
[8 empty] | 1.991 | g
[8 empty] | 1.997 | g
[8 empty] | 1.99 | g
[8 empty] | 2.004 | g
[8 empty] | 1.993 | g
[8 empty] | 1.996 | g
[8 empty] | 1.998 | g
[8 empty] | 1.997 | g
[8 empty] | 2.001 | g
[8 empty] | 1.993 | g
[8 empty] | 1.993 | g
[8 empty] | 1.986 | g
[8 empty] | 1.989 | g
[8 empty] | 1.985 | g
[8 empty] | 1.991 | g
[8 empty] | 1.991 | g
[8 empty] | 1.995 | g
[8 empty] | 1.996 | g
[8 empty] | 1.991 | g
[8 empty] | 1.989 | g
[8 empty] | 1.993 | g
[8 empty] | 1.996 | g
[8 empty] | 1.994 | g
[8 empty] | 1.985 | g
[8 empty] | 1.271 | g
[8 empty] | -0.015 | g
[8 empty] | -0.007 | g
[8 empty] | -0.007 | g
[8 empty] | -0.008 | g
[8 empty] | -0.011 | g
[8 empty] | -0.018 | g
[8 empty] | -0.006 | g
[8 empty] | -0.005 | g
[8 empty] | -0.003 | g
[8 empty] | -0.002 | g
[8 empty] | 0.001 | g
[8 empty] | 3.145 | g
[8 empty] | 5.001 | g
[8 empty] | 4.995 | g
[8 empty] | 4.993 | g
[8 empty] | 4.987 | g
[8 empty] | 4.988 | g
[8 empty] | 4.994 | g
[8 empty] | 4.989 | g
[8 empty] | 4.983 | g
[8 empty] | 4.985 | g
[8 empty] | 4.987 | g
[8 empty] | 4.984 | g
[8 empty] | 4.992 | g
[8 empty] | 4.987 | g
[8 empty] | 4.984 | g
[8 empty] | 4.984 | g
[8 empty] | 4.985 | g
[8 empty] | 4.995 | g
[8 empty] | 4.978 | g
[8 empty] | 4.984 | g
[8 empty] | 4.985 | g
[8 empty] | 4.987 | g
[8 empty] | 4.985 | g
[8 empty] | 4.982 | g
[8 empty] | 4.986 | g
[8 empty] | 4.983 | g
[8 empty] | 4.988 | g
[8 empty] | 4.987 | g
[8 empty] | 4.985 | g
[8 empty] | 4.981 | g
[8 empty] | 4.985 | g
[8 empty] | 4.981 | g
[8 empty] | 2.46 | g
[8 empty] | -0.01 | g
[8 empty] | -0.009 | g
[8 empty] | -0.019 | g
[8 empty] | 0.002 | g
[8 empty] | -0.005 | g
[8 empty] | -0.014 | g
[8 empty] | -0.003 | g
[8 empty] | -0.009 | g
[8 empty] | -0.009 | g
[8 empty] | -0.006 | g
[8 empty] | -0.011 | g
[8 empty] | -0.009 | g
[8 empty] | 7.231 | g
[8 empty] | 9.988 | g
[8 empty] | 9.991 | g
[8 empty] | 9.987 | g
[8 empty] | 9.979 | g
[8 empty] | 9.983 | g
[8 empty] | 9.996 | g
[8 empty] | 9.984 | g
[8 empty] | 9.99 | g
[8 empty] | 9.988 | g
[8 empty] | 9.998 | g
[8 empty] | 9.991 | g
[8 empty] | 9.988 | g
[8 empty] | 9.997 | g
[8 empty] | 9.988 | g
[8 empty] | 9.996 | g
[8 empty] | 9.992 | g
[8 empty] | 9.993 | g
[8 empty] | 9.989 | g
[8 empty] | 9.997 | g
[8 empty] | 9.999 | g
[8 empty] | 9.991 | g
[8 empty] | 9.993 | g
[8 empty] | 9.991 | g
[8 empty] | 9.991 | g
[8 empty] | 9.985 | g
[8 empty] | 9.99 | g
[8 empty] | 9.999 | g
[8 empty] | 10.008 | g
[8 empty] | 2.477 | g
[8 empty] | -0.004 | g
[8 empty] | -0.002 | g
[8 empty] | -0.005 | g
[8 empty] | -0.003 | g
[8 empty] | -0.015 | g
[8 empty] | -0.005 | g
[8 empty] | 0.005 | g
[8 empty] | -0.004 | g
[8 empty] | 19.71 | g
[8 empty] | 19.99 | g
[8 empty] | 19.987 | g
[8 empty] | 19.985 | g
[8 empty] | 19.993 | g
[8 empty] | 19.987 | g
[8 empty] | 19.991 | g
[8 empty] | 19.982 | g
[8 empty] | 19.985 | g
[8 empty] | 19.979 | g
[8 empty] | 19.981 | g
[8 empty] | 19.993 | g
[8 empty] | 19.991 | g
[8 empty] | 19.994 | g
[8 empty] | 19.995 | g
[8 empty] | 19.978 | g
[8 empty] | 19.994 | g
[8 empty] | 19.987 | g
[8 empty] | 19.998 | g
[8 empty] | 19.992 | g
[8 empty] | 19.986 | g
[8 empty] | 19.993 | g
[8 empty] | 19.989 | g
[8 empty] | 19.995 | g
[8 empty] | 19.997 | g
[8 empty] | 19.99 | g
[8 empty] | 19.99 | g
[8 empty] | 16.456 | g
[8 empty] | -0.007 | g
[8 empty] | -0.007 | g
[8 empty] | -0.005 | g
[8 empty] | 0 | g
[8 empty] | -0.005 | g
[8 empty] | -0.002 | g
[8 empty] | -0.006 | g
[8 empty] | 34.341 | g
[8 empty] | 49.984 | g
[8 empty] | 49.984 | g
[8 empty] | 49.988 | g
[8 empty] | 49.983 | g
[8 empty] | 49.988 | g
[8 empty] | 49.982 | g
[8 empty] | 49.983 | g
[8 empty] | 49.99 | g
[8 empty] | 49.988 | g
[8 empty] | 49.987 | g
[8 empty] | 49.99 | g
[8 empty] | 49.989 | g
[8 empty] | 49.993 | g
[8 empty] | 49.984 | g
[8 empty] | 49.994 | g
[8 empty] | 50 | g
[8 empty] | 49.993 | g
[8 empty] | 49.999 | g
[8 empty] | 49.999 | g
[8 empty] | 50.009 | g
[8 empty] | 50.003 | g
[8 empty] | 50.004 | g
[8 empty] | 49.998 | g
[8 empty] | 49.993 | g
[8 empty] | 49.999 | g
[8 empty] | 49.992 | g
[8 empty] | 53.558 | g
[8 empty] | 8.672 | g
[8 empty] | 0.001 | g
[8 empty] | 0.002 | g
[8 empty] | 0.006 | g
[8 empty] | 0.007 | g
[8 empty] | 0 | g
[8 empty] | -0.002 | g
[8 empty] | -0.006 | g
[8 empty] | 0.008 | g
[8 empty] | -0.008 | g
[8 empty] | 0.005 | g
[8 empty] | -0.004 | g
[8 empty] | 0.001 | g
[8 empty] | 71.066 | g
[8 empty] | 99.937 | g
[8 empty] | 99.949 | g
[8 empty] | 99.954 | g
[8 empty] | 99.95 | g
[8 empty] | 99.959 | g
[8 empty] | 99.953 | g
[8 empty] | 99.955 | g
[8 empty] | 99.952 | g
[8 empty] | 99.955 | g
[8 empty] | 99.964 | g
[8 empty] | 99.957 | g
[8 empty] | 99.961 | g
[8 empty] | 99.964 | g
[8 empty] | 99.97 | g
[8 empty] | 99.967 | g
[8 empty] | 99.979 | g
[8 empty] | 99.975 | g
[8 empty] | 99.982 | g
[8 empty] | 99.988 | g
[8 empty] | 99.982 | g
[8 empty] | 99.986 | g
[8 empty] | 99.984 | g
[8 empty] | 99.994 | g
[8 empty] | 99.984 | g
[8 empty] | 99.987 | g
[8 empty] | 99.989 | g
[8 empty] | 86.839 | g
[8 empty] | 0.039 | g
[8 empty] | 0.019 | g
[8 empty] | 0.027 | g
[8 empty] | 0.023 | g
[8 empty] | 0.022 | g
[8 empty] | 0.02 | g
[8 empty] | 0.014 | g
[8 empty] | 0.014 | g
[8 empty] | 0.024 | g
[8 empty] | 0.013 | g
[8 empty] | 0.01 | g
[8 empty] | 0.001 | g
[8 empty] | 0.008 | g
[8 empty] | 0.006 | g
[8 empty] | 0.004 | g
[8 empty] | 0.012 | g
[8 empty] | 0.01 | g
[8 empty] | 0 | g
[8 empty] | 0.016 | g
[8 empty] | 0.013 | g
[8 empty] | 0.012 | g
[8 empty] | 0.003 | g
[8 empty] | -0.006 | g
[8 empty] | 0.005 | g

## 200623 - 5000g Standards PLA
Mass Balance Test
Device Under Test: | OS_Nano_Balance (3D printed balance), Using TAL220 5kg load cell
Lab Balance: | N/A Using standard masses
DUT Calibration: | Calibrated with 100 g standard mass
| 424.094 | g/div
PuTTY
[9 empty] | -0.01038 | g
|  |  |  | Averages are corrected for zero offset (from drift) |  |  |  |  | -0.00684 | g
|  |  | Nominal | Average | Std | Count | Addresses | Zero offset from drift | -0.01486 | g
|  |  | 1 | 1.0083 | 0.029811 | 33 | R42:R74 | 0.0083421 | -0.04056 | g
|  |  | 2 | 2.0534 | 0.047071 | 28 | R87:R114 | 0.02916 | -0.04409 | g
|  |  | 5 | 4.9196 | 0.043586 | 28 | R121:R148 | 0.15003 | 0.01038 | g
|  |  | 10 | 10.0645 | 0.034261 | 32 | R158:R189 | 0.05626 | -0.01627 | g
|  |  | 20 | 19.994 | 0.03144 | 41 | R199:R239 | 0.12356 | -0.0132 | g
|  |  | 50 | 50.0201 | 0.031772 | 34 | R251:R284 | 0.13355 | -0.02334 | g
|  |  | 100 | 99.9603 | 0.038972 | 32 | R297:R328 | 0.14067 | 0.04905 | g
|  |  | 1 | 0.98272 | 0.040242 | 37 | R342:R378 | 0.23457 | -0.0573 | g
|  |  | 2 | 2.0255 | 0.03306 | 29 | R394:R422 | 0.18505 | -0.01792 | g
|  |  | 5 | 4.99 | 0.034046 | 37 | R432:R468 | 0.23789 | 0.01038 | g
|  |  | 10 | 10.0106 | 0.034037 | 28 | R478:R505 | 0.19684 | -0.01368 | g
|  |  | 20 | 19.9614 | 0.045823 | 26 | R516:R541 | 0.28838 | 0.00637 | g
|  |  | 50 | 50.0077 | 0.040101 | 33 | R555:R587 | 0.2249 | 0.02452 | g
|  |  | 100 | 99.9869 | 0.027839 | 30 | R604:R633 | 0.25169 | -0.01014 | g
[9 empty] | 0.04622 | g
|  |  |  | Avg | 0.036576 |  | 0.0007315% |  | 0.03254 | g
|  |  |  | Max | 0.047071 |  |  |  | 0.04221 | g
[9 empty] | -0.0481 | g
[9 empty] | 0.07946 | g
[9 empty] | -0.02476 | g
[9 empty] | 0.06956 | g
[9 empty] | 0.05918 | g
[9 empty] | 0.06885 | g
[9 empty] | 0.01886 | g
[9 empty] | 0.05423 | g
[9 empty] | 0.01155 | g
[9 empty] | 0.15963 | g
[9 empty] | 1.361 | g
[9 empty] | 1.0238 | g
[9 empty] | 0.98068 | g
[9 empty] | 1.0387 | g
[9 empty] | 1.0102 | g
[9 empty] | 1.038 | g
[9 empty] | 1.0592 | g
[9 empty] | 1.004 | g
[9 empty] | 0.98705 | g
[9 empty] | 1.0196 | g
[9 empty] | 1.016 | g
[9 empty] | 1.0003 | g
[9 empty] | 0.96441 | g
[9 empty] | 1.0311 | g
[9 empty] | 1.0017 | g
[9 empty] | 1.0215 | g
[9 empty] | 1.0283 | g
[9 empty] | 1.084 | g
[9 empty] | 1.0691 | g
[9 empty] | 1.0113 | g
[9 empty] | 0.97408 | g
[9 empty] | 1.0488 | g
[9 empty] | 0.99577 | g
[9 empty] | 0.99058 | g
[9 empty] | 0.97455 | g
[9 empty] | 1.0123 | g
[9 empty] | 1.0196 | g
[9 empty] | 0.9736 | g
[9 empty] | 1.0333 | g
[9 empty] | 0.9795 | g
[9 empty] | 1.0625 | g
[9 empty] | 1.0594 | g
[9 empty] | 1.0219 | g
[9 empty] | 1.0137 | g
[9 empty] | 0.249 | g
[9 empty] | 0.02665 | g
[9 empty] | 0.0797 | g
[9 empty] | -0.03726 | g
[9 empty] | -0.05022 | g
[9 empty] | 0.00495 | g
[9 empty] | -0.01368 | g
[9 empty] | 0.0323 | g
[9 empty] | -0.029 | g
[9 empty] | 0.02688 | g
[9 empty] | 0.52936 | g
[9 empty] | 2.051 | g
[9 empty] | 2.0043 | g
[9 empty] | 2.021 | g
[9 empty] | 2.0337 | g
[9 empty] | 2.0724 | g
[9 empty] | 2.0503 | g
[9 empty] | 2.0208 | g
[9 empty] | 2.1184 | g
[9 empty] | 1.9833 | g
[9 empty] | 2.0356 | g
[9 empty] | 2.0849 | g
[9 empty] | 2.0748 | g
[9 empty] | 2.1399 | g
[9 empty] | 2.142 | g
[9 empty] | 2.1243 | g
[9 empty] | 2.0493 | g
[9 empty] | 2.0781 | g
[9 empty] | 2.1538 | g
[9 empty] | 2.1311 | g
[9 empty] | 2.1354 | g
[9 empty] | 2.116 | g
[9 empty] | 2.0672 | g
[9 empty] | 2.0951 | g
[9 empty] | 2.0936 | g
[9 empty] | 2.0974 | g
[9 empty] | 2.0811 | g
[9 empty] | 2.034 | g
[9 empty] | 2.1632 | g
[9 empty] | 2.1099 | g
[9 empty] | 1.0234 | g
[9 empty] | 0.19783 | g
[9 empty] | 0.09927 | g
[9 empty] | 0.18911 | g
[9 empty] | 0.11389 | g
[9 empty] | 0.42538 | g
[9 empty] | 5.0501 | g
[9 empty] | 5.0788 | g
[9 empty] | 5.1076 | g
[9 empty] | 5.0399 | g
[9 empty] | 5.093 | g
[9 empty] | 5.084 | g
[9 empty] | 5.0557 | g
[9 empty] | 5.1262 | g
[9 empty] | 5.1368 | g
[9 empty] | 5.1196 | g
[9 empty] | 5.0984 | g
[9 empty] | 5.0623 | g
[9 empty] | 5.0416 | g
[9 empty] | 5.063 | g
[9 empty] | 5.0362 | g
[9 empty] | 5.0607 | g
[9 empty] | 5.0489 | g
[9 empty] | 5.0543 | g
[9 empty] | 5.0196 | g
[9 empty] | 4.9322 | g
[9 empty] | 5.0609 | g
[9 empty] | 5.0984 | g
[9 empty] | 5.1474 | g
[9 empty] | 5.0781 | g
[9 empty] | 5.1067 | g
[9 empty] | 5.105 | g
[9 empty] | 5.0281 | g
[9 empty] | 5.0159 | g
[9 empty] | 5.1272 | g
[9 empty] | 5.1562 | g
[9 empty] | 3.8798 | g
[9 empty] | 0.07545 | g
[9 empty] | 0.07663 | g
[9 empty] | 0.05022 | g
[9 empty] | 0.0316 | g
[9 empty] | 0.0474 | g
[9 empty] | 9.8488 | g
[9 empty] | 10.1244 | g
[9 empty] | 10.1282 | g
[9 empty] | 10.1253 | g
[9 empty] | 10.0945 | g
[9 empty] | 10.086 | g
[9 empty] | 10.0487 | g
[9 empty] | 10.119 | g
[9 empty] | 10.1548 | g
[9 empty] | 10.1338 | g
[9 empty] | 10.1128 | g
[9 empty] | 10.0704 | g
[9 empty] | 10.1395 | g
[9 empty] | 10.1195 | g
[9 empty] | 10.164 | g
[9 empty] | 10.1619 | g
[9 empty] | 10.0992 | g
[9 empty] | 10.1202 | g
[9 empty] | 10.1272 | g
[9 empty] | 10.0846 | g
[9 empty] | 10.0921 | g
[9 empty] | 10.0949 | g
[9 empty] | 10.0501 | g
[9 empty] | 10.1253 | g
[9 empty] | 10.1654 | g
[9 empty] | 10.1447 | g
[9 empty] | 10.165 | g
[9 empty] | 10.0959 | g
[9 empty] | 10.119 | g
[9 empty] | 10.1831 | g
[9 empty] | 10.0999 | g
[9 empty] | 10.1256 | g
[9 empty] | 10.1895 | g
[9 empty] | 10.0801 | g
[9 empty] | 10.1204 | g
[9 empty] | 1.4244 | g
[9 empty] | 0.20679 | g
[9 empty] | 0.20231 | g
[9 empty] | 0.03419 | g
[9 empty] | 0.08324 | g
[9 empty] | 0.09125 | g
[9 empty] | 18.5735 | g
[9 empty] | 20.079 | g
[9 empty] | 20.0762 | g
[9 empty] | 20.0838 | g
[9 empty] | 20.0885 | g
[9 empty] | 20.0437 | g
[9 empty] | 20.0906 | g
[9 empty] | 20.1453 | g
[9 empty] | 20.1295 | g
[9 empty] | 20.1264 | g
[9 empty] | 20.1526 | g
[9 empty] | 20.147 | g
[9 empty] | 20.1182 | g
[9 empty] | 20.1172 | g
[9 empty] | 20.1205 | g
[9 empty] | 20.0913 | g
[9 empty] | 20.084 | g
[9 empty] | 20.1708 | g
[9 empty] | 20.104 | g
[9 empty] | 20.1208 | g
[9 empty] | 20.1215 | g
[9 empty] | 20.1264 | g
[9 empty] | 20.1088 | g
[9 empty] | 20.1132 | g
[9 empty] | 20.142 | g
[9 empty] | 20.0991 | g
[9 empty] | 20.1194 | g
[9 empty] | 20.0604 | g
[9 empty] | 20.1576 | g
[9 empty] | 20.1708 | g
[9 empty] | 20.0906 | g
[9 empty] | 20.0979 | g
[9 empty] | 20.1842 | g
[9 empty] | 20.1528 | g
[9 empty] | 20.1135 | g
[9 empty] | 20.1043 | g
[9 empty] | 20.1399 | g
[9 empty] | 20.1208 | g
[9 empty] | 20.0753 | g
[9 empty] | 20.1205 | g
[9 empty] | 20.1639 | g
[9 empty] | 20.1484 | g
[9 empty] | 15.7194 | g
[9 empty] | 0.12474 | g
[9 empty] | 0.11601 | g
[9 empty] | 0.15563 | g
[9 empty] | 0.11342 | g
[9 empty] | 0.13582 | g
[9 empty] | 0.13252 | g
[9 empty] | 0.12615 | g
[9 empty] | 0.16411 | g
[9 empty] | 6.0267 | g
[9 empty] | 50.4157 | g
[9 empty] | 50.2049 | g
[9 empty] | 50.1703 | g
[9 empty] | 50.1875 | g
[9 empty] | 50.1294 | g
[9 empty] | 50.129 | g
[9 empty] | 50.179 | g
[9 empty] | 50.1809 | g
[9 empty] | 50.1349 | g
[9 empty] | 50.1372 | g
[9 empty] | 50.1377 | g
[9 empty] | 50.12 | g
[9 empty] | 50.1184 | g
[9 empty] | 50.1486 | g
[9 empty] | 50.1759 | g
[9 empty] | 50.0941 | g
[9 empty] | 50.154 | g
[9 empty] | 50.1594 | g
[9 empty] | 50.1099 | g
[9 empty] | 50.1009 | g
[9 empty] | 50.1771 | g
[9 empty] | 50.1594 | g
[9 empty] | 50.179 | g
[9 empty] | 50.1627 | g
[9 empty] | 50.166 | g
[9 empty] | 50.1195 | g
[9 empty] | 50.1118 | g
[9 empty] | 50.1769 | g
[9 empty] | 50.174 | g
[9 empty] | 50.1988 | g
[9 empty] | 50.1236 | g
[9 empty] | 50.1448 | g
[9 empty] | 50.2212 | g
[9 empty] | 50.2023 | g
[9 empty] | 50.1344 | g
[9 empty] | 22.088 | g
[9 empty] | 0.12379 | g
[9 empty] | 0.18769 | g
[9 empty] | 0.13629 | g
[9 empty] | 0.19901 | g
[9 empty] | 0.13228 | g
[9 empty] | 0.17025 | g
[9 empty] | 0.10328 | g
[9 empty] | 0.09668 | g
[9 empty] | 0.11672 | g
[9 empty] | 19.7697 | g
[9 empty] | 100.2643 | g
[9 empty] | 100.0436 | g
[9 empty] | 100.1684 | g
[9 empty] | 100.1238 | g
[9 empty] | 100.082 | g
[9 empty] | 100.0658 | g
[9 empty] | 100.0995 | g
[9 empty] | 100.1033 | g
[9 empty] | 100.1054 | g
[9 empty] | 100.0408 | g
[9 empty] | 100.0974 | g
[9 empty] | 100.0915 | g
[9 empty] | 100.0837 | g
[9 empty] | 100.1 | g
[9 empty] | 100.1007 | g
[9 empty] | 100.0615 | g
[9 empty] | 100.1372 | g
[9 empty] | 100.0696 | g
[9 empty] | 100.0851 | g
[9 empty] | 100.1172 | g
[9 empty] | 100.1441 | g
[9 empty] | 100.0571 | g
[9 empty] | 100.0854 | g
[9 empty] | 100.0939 | g
[9 empty] | 100.0766 | g
[9 empty] | 100.1686 | g
[9 empty] | 100.117 | g
[9 empty] | 100.074 | g
[9 empty] | 100.1401 | g
[9 empty] | 100.0679 | g
[9 empty] | 100.2285 | g
[9 empty] | 100.1243 | g
[9 empty] | 100.0769 | g
[9 empty] | 97.3996 | g
[9 empty] | 0.40651 | g
[9 empty] | 0.18934 | g
[9 empty] | 0.27069 | g
[9 empty] | 0.20962 | g
[9 empty] | 0.27989 | g
[9 empty] | 0.26268 | g
[9 empty] | 0.19312 | g
[9 empty] | 0.2299 | g
[9 empty] | 0.23933 | g
[9 empty] | 0.23933 | g
[9 empty] | 0.23179 | g
[9 empty] | 0.87363 | g
[9 empty] | 1.2443 | g
[9 empty] | 1.1356 | g
[9 empty] | 1.2542 | g
[9 empty] | 1.2273 | g
[9 empty] | 1.2405 | g
[9 empty] | 1.2233 | g
[9 empty] | 1.1995 | g
[9 empty] | 1.2558 | g
[9 empty] | 1.1736 | g
[9 empty] | 1.2658 | g
[9 empty] | 1.2287 | g
[9 empty] | 1.228 | g
[9 empty] | 1.1969 | g
[9 empty] | 1.2676 | g
[9 empty] | 1.1745 | g
[9 empty] | 1.2518 | g
[9 empty] | 1.2052 | g
[9 empty] | 1.2377 | g
[9 empty] | 1.2804 | g
[9 empty] | 1.1936 | g
[9 empty] | 1.2431 | g
[9 empty] | 1.183 | g
[9 empty] | 1.274 | g
[9 empty] | 1.2257 | g
[9 empty] | 1.2259 | g
[9 empty] | 1.1868 | g
[9 empty] | 1.2783 | g
[9 empty] | 1.2438 | g
[9 empty] | 1.207 | g
[9 empty] | 1.2028 | g
[9 empty] | 1.2059 | g
[9 empty] | 1.2082 | g
[9 empty] | 1.1052 | g
[9 empty] | 1.1391 | g
[9 empty] | 1.1637 | g
[9 empty] | 1.2361 | g
[9 empty] | 1.2268 | g
[9 empty] | 1.5447 | g
[9 empty] | 0.19642 | g
[9 empty] | 0.16624 | g
[9 empty] | 0.20915 | g
[9 empty] | 0.17803 | g
[9 empty] | 0.19147 | g
[9 empty] | 0.1436 | g
[9 empty] | 0.22141 | g
[9 empty] | 0.21976 | g
[9 empty] | 0.18486 | g
[9 empty] | 0.13275 | g
[9 empty] | 0.22566 | g
[9 empty] | 0.16246 | g
[9 empty] | 0.17378 | g
[9 empty] | 2.4554 | g
[9 empty] | 2.2561 | g
[9 empty] | 2.183 | g
[9 empty] | 2.1875 | g
[9 empty] | 2.1644 | g
[9 empty] | 2.1517 | g
[9 empty] | 2.1722 | g
[9 empty] | 2.2403 | g
[9 empty] | 2.1705 | g
[9 empty] | 2.2247 | g
[9 empty] | 2.1929 | g
[9 empty] | 2.2045 | g
[9 empty] | 2.1891 | g
[9 empty] | 2.2021 | g
[9 empty] | 2.2054 | g
[9 empty] | 2.1745 | g
[9 empty] | 2.2726 | g
[9 empty] | 2.1599 | g
[9 empty] | 2.267 | g
[9 empty] | 2.1986 | g
[9 empty] | 2.2266 | g
[9 empty] | 2.225 | g
[9 empty] | 2.2297 | g
[9 empty] | 2.2493 | g
[9 empty] | 2.2075 | g
[9 empty] | 2.2434 | g
[9 empty] | 2.253 | g
[9 empty] | 2.195 | g
[9 empty] | 2.2443 | g
[9 empty] | 2.2148 | g
[9 empty] | 2.287 | g
[9 empty] | 2.2137 | g
[9 empty] | 0.22354 | g
[9 empty] | 0.28461 | g
[9 empty] | 0.22212 | g
[9 empty] | 0.23462 | g
[9 empty] | 0.24924 | g
[9 empty] | 0.24098 | g
[9 empty] | 0.21009 | g
[9 empty] | 5.2142 | g
[9 empty] | 5.2184 | g
[9 empty] | 5.1663 | g
[9 empty] | 5.2092 | g
[9 empty] | 5.229 | g
[9 empty] | 5.2241 | g
[9 empty] | 5.1951 | g
[9 empty] | 5.1706 | g
[9 empty] | 5.2111 | g
[9 empty] | 5.2385 | g
[9 empty] | 5.1607 | g
[9 empty] | 5.2599 | g
[9 empty] | 5.251 | g
[9 empty] | 5.217 | g
[9 empty] | 5.2054 | g
[9 empty] | 5.1897 | g
[9 empty] | 5.2444 | g
[9 empty] | 5.2632 | g
[9 empty] | 5.2382 | g
[9 empty] | 5.3201 | g
[9 empty] | 5.2484 | g
[9 empty] | 5.247 | g
[9 empty] | 5.2748 | g
[9 empty] | 5.2194 | g
[9 empty] | 5.2158 | g
[9 empty] | 5.2804 | g
[9 empty] | 5.2243 | g
[9 empty] | 5.2076 | g
[9 empty] | 5.2253 | g
[9 empty] | 5.226 | g
[9 empty] | 5.1885 | g
[9 empty] | 5.2844 | g
[9 empty] | 5.2238 | g
[9 empty] | 5.2689 | g
[9 empty] | 5.1986 | g
[9 empty] | 5.2642 | g
[9 empty] | 5.208 | g
[9 empty] | 5.1514 | g
[9 empty] | 0.26669 | g
[9 empty] | 0.17968 | g
[9 empty] | 0.23745 | g
[9 empty] | 0.16176 | g
[9 empty] | 0.19689 | g
[9 empty] | 0.20844 | g
[9 empty] | 1.1806 | g
[9 empty] | 10.1944 | g
[9 empty] | 10.1824 | g
[9 empty] | 10.1624 | g
[9 empty] | 10.2506 | g
[9 empty] | 10.2041 | g
[9 empty] | 10.1954 | g
[9 empty] | 10.156 | g
[9 empty] | 10.2284 | g
[9 empty] | 10.2432 | g
[9 empty] | 10.1966 | g
[9 empty] | 10.2107 | g
[9 empty] | 10.1923 | g
[9 empty] | 10.2468 | g
[9 empty] | 10.1907 | g
[9 empty] | 10.205 | g
[9 empty] | 10.1461 | g
[9 empty] | 10.2119 | g
[9 empty] | 10.2388 | g
[9 empty] | 10.2145 | g
[9 empty] | 10.1789 | g
[9 empty] | 10.2013 | g
[9 empty] | 10.2223 | g
[9 empty] | 10.1699 | g
[9 empty] | 10.1883 | g
[9 empty] | 10.2376 | g
[9 empty] | 10.2732 | g
[9 empty] | 10.2569 | g
[9 empty] | 10.256 | g
[9 empty] | 10.1482 | g
[9 empty] | 10.2048 | g
[9 empty] | 3.9121 | g
[9 empty] | 0.32422 | g
[9 empty] | 0.22542 | g
[9 empty] | 0.31385 | g
[9 empty] | 0.32681 | g
[9 empty] | 0.27612 | g
[9 empty] | 0.26386 | g
[9 empty] | 6.2967 | g
[9 empty] | 20.2299 | g
[9 empty] | 20.2948 | g
[9 empty] | 20.2146 | g
[9 empty] | 20.2568 | g
[9 empty] | 20.3014 | g
[9 empty] | 20.2288 | g
[9 empty] | 20.2908 | g
[9 empty] | 20.3054 | g
[9 empty] | 20.2788 | g
[9 empty] | 20.2797 | g
[9 empty] | 20.2531 | g
[9 empty] | 20.2696 | g
[9 empty] | 20.2203 | g
[9 empty] | 20.1925 | g
[9 empty] | 20.321 | g
[9 empty] | 20.1972 | g
[9 empty] | 20.3198 | g
[9 empty] | 20.3195 | g
[9 empty] | 20.2083 | g
[9 empty] | 20.2149 | g
[9 empty] | 20.212 | g
[9 empty] | 20.2672 | g
[9 empty] | 20.242 | g
[9 empty] | 20.158 | g
[9 empty] | 20.2422 | g
[9 empty] | 20.1722 | g
[9 empty] | 20.2333 | g
[9 empty] | 20.2519 | g
[9 empty] | 12.1822 | g
[9 empty] | 0.17991 | g
[9 empty] | 0.23981 | g
[9 empty] | 0.26787 | g
[9 empty] | 0.23226 | g
[9 empty] | 0.26504 | g
[9 empty] | 0.2108 | g
[9 empty] | 0.17378 | g
[9 empty] | 0.22471 | g
[9 empty] | 0.2299 | g
[9 empty] | 16.5815 | g
[9 empty] | 50.1893 | g
[9 empty] | 50.1052 | g
[9 empty] | 50.1653 | g
[9 empty] | 50.1978 | g
[9 empty] | 50.2308 | g
[9 empty] | 50.1842 | g
[9 empty] | 50.1974 | g
[9 empty] | 50.2155 | g
[9 empty] | 50.2485 | g
[9 empty] | 50.1823 | g
[9 empty] | 50.2495 | g
[9 empty] | 50.2264 | g
[9 empty] | 50.2884 | g
[9 empty] | 50.2011 | g
[9 empty] | 50.2497 | g
[9 empty] | 50.2344 | g
[9 empty] | 50.232 | g
[9 empty] | 50.2738 | g
[9 empty] | 50.2886 | g
[9 empty] | 50.2209 | g
[9 empty] | 50.1957 | g
[9 empty] | 50.2752 | g
[9 empty] | 50.2742 | g
[9 empty] | 50.2749 | g
[9 empty] | 50.2334 | g
[9 empty] | 50.2436 | g
[9 empty] | 50.1995 | g
[9 empty] | 50.2551 | g
[9 empty] | 50.2464 | g
[9 empty] | 50.2136 | g
[9 empty] | 50.2856 | g
[9 empty] | 50.27 | g
[9 empty] | 50.2457 | g
[9 empty] | 50.2698 | g
[9 empty] | 14.3188 | g
[9 empty] | 0.25867 | g
[9 empty] | 0.24193 | g
[9 empty] | 0.25678 | g
[9 empty] | 0.29003 | g
[9 empty] | 0.26881 | g
[9 empty] | 0.2075 | g
[9 empty] | 0.25867 | g
[9 empty] | 0.29828 | g
[9 empty] | 0.20066 | g
[9 empty] | 0.24098 | g
[9 empty] | 0.25725 | g
[9 empty] | 0.24075 | g
[9 empty] | 1.6593 | g
[9 empty] | 97.0033 | g
[9 empty] | 100.1587 | g
[9 empty] | 100.2448 | g
[9 empty] | 100.2002 | g
[9 empty] | 100.254 | g
[9 empty] | 100.2695 | g
[9 empty] | 100.1754 | g
[9 empty] | 100.2445 | g
[9 empty] | 100.2099 | g
[9 empty] | 100.1971 | g
[9 empty] | 100.2448 | g
[9 empty] | 100.1747 | g
[9 empty] | 100.2667 | g
[9 empty] | 100.2441 | g
[9 empty] | 100.2273 | g
[9 empty] | 100.2469 | g
[9 empty] | 100.266 | g
[9 empty] | 100.2278 | g
[9 empty] | 100.2179 | g
[9 empty] | 100.253 | g
[9 empty] | 100.2155 | g
[9 empty] | 100.295 | g
[9 empty] | 100.2193 | g
[9 empty] | 100.2434 | g
[9 empty] | 100.2756 | g
[9 empty] | 100.2523 | g
[9 empty] | 100.2257 | g
[9 empty] | 100.2523 | g
[9 empty] | 100.258 | g
[9 empty] | 100.2681 | g
[9 empty] | 100.2457 | g
[9 empty] | 100.2431 | g
[9 empty] | 100.2926 | g
[9 empty] | 67.3603 | g
[9 empty] | 0.26268 | g
[9 empty] | 0.31691 | g
[9 empty] | 0.25348 | g
[9 empty] | 0.28484 | g
[9 empty] | 0.27753 | g
[9 empty] | 0.25678 | g
[9 empty] | 0.29805 | g
[9 empty] | 0.22872 | g
[9 empty] | 0.29003 | g
[9 empty] | 0.21528 | g
[9 empty] | 0.18958 | g
[9 empty] | 0.32799 | g
[9 empty] | 0.31809 | g
[9 empty] | 0.25254 | g
[9 empty] | 0.32257 | g
[9 empty] | 0.23179 | g
[9 empty] | 0.26338 | g
[9 empty] | 0.2549 | g
[9 empty] | 0.26716 | g
[9 empty] | 0.28508 | g
[9 empty] | 0.2872 | g
[9 empty] | 0.24193 | g
[9 empty] | 0.31314 | g

## 200624 - 5000g Standards Wood
Mass Balance Test
Device Under Test: | OS_Nano_Balance (3D printed balance), Using TAL220 5kg load cell on wood frame
Lab Balance: | N/A Using standard masses
DUT Calibration: | Calibrated with 100 g standard mass
| 424.094 | g/div
PuTTY
[9 empty] | -0.01674 | g
[9 empty] | 0.00236 | g
|  |  | Nominal | Average | Std | Count | Addresses |  | -0.07616 | g
|  |  | 1 | 0.99919 | 0.036326 | 31 | R66:R96 | -0.047563 | -0.02476 | g
|  |  | 2 | 2.0167 | 0.038287 | 32 | R106:R137 | -0.048507 | -0.08041 | g
|  |  | 5 | 4.9826 | 0.028529 | 36 | R161:R196 | -0.0297 | -0.04032 | g
|  |  | 10 | 9.9652 | 0.035934 | 31 | R206:R236 | -0.007625 | -0.0158 | g
|  |  | 20 | 20.0059 | 0.03411 | 34 | R248:R281 | -0.04551 | 0.00071 | g
|  |  | 50 | 49.9839 | 0.036011 | 37 | R293:R329 | -0.027757 | -0.04999 | g
|  |  | 100 | 99.9872 | 0.030031 | 30 | R345:R374 | -0.038606 | -0.14643 | g
|  |  | 1 | 1.0079 | 0.032765 | 32 | R389:R420 | -0.021293 | -0.03773 | g
|  |  | 2 | 2.0259 | 0.03416 | 31 | R431:R461 | -0.03258 | -0.02806 | g
|  |  | 5 | 4.9999 | 0.036244 | 33 | R472:R504 | -0.023849 | -0.07687 | g
|  |  | 10 | 10.0099 | 0.027428 | 45 | R516:R560 | -0.028533 | -0.1021 | g
|  |  | 20 | 19.9942 | 0.036041 | 38 | R570:R607 | -0.021788 | -0.01721 | g
|  |  | 50 | 49.9715 | 0.030486 | 42 | R622:R663 | -0.0075711 | -0.01768 | g
|  |  | 100 | 99.9749 | 0.034404 | 37 | R679:R715 | -0.019677 | -0.09998 | g
[9 empty] | -0.01698 | g
|  |  |  | Avg | 0.033626 |  |  |  | -0.06154 | g
|  |  |  | Max | 0.038287 |  |  |  | -0.04032 | g
[9 empty] | -0.06673 | g
[9 empty] | -0.05164 | g
[9 empty] | -0.0191 | g
[9 empty] | -0.03018 | g
[9 empty] | -0.02971 | g
[9 empty] | -0.08041 | g
[9 empty] | -0.03891 | g
[9 empty] | -0.00047 | g
[9 empty] | -0.04433 | g
[9 empty] | -0.10493 | g
[9 empty] | -0.00495 | g
[9 empty] | 0.00448 | g
[9 empty] | -0.08017 | g
[9 empty] | -0.10446 | g
[9 empty] | -0.03089 | g
[9 empty] | 0.00637 | g
[9 empty] | -0.00165 | g
[9 empty] | -0.1344 | g
[9 empty] | -0.06437 | g
[9 empty] | -0.03466 | g
[9 empty] | 0.03678 | g
[9 empty] | -0.07593 | g
[9 empty] | -0.04669 | g
[9 empty] | -0.07734 | g
[9 empty] | 0.01863 | g
[9 empty] | -0.04315 | g
[9 empty] | 0.00731 | g
[9 empty] | -0.06201 | g
[9 empty] | -0.10847 | g
[9 empty] | -0.07121 | g
[9 empty] | -0.02358 | g
[9 empty] | -0.06909 | g
[9 empty] | -0.03985 | g
[9 empty] | 0.65363 | g
[9 empty] | 0.9505 | g
[9 empty] | 1.0146 | g
[9 empty] | 0.89202 | g
[9 empty] | 1.0342 | g
[9 empty] | 0.89603 | g
[9 empty] | 0.93541 | g
[9 empty] | 0.93258 | g
[9 empty] | 0.90522 | g
[9 empty] | 0.94437 | g
[9 empty] | 0.98091 | g
[9 empty] | 1.0186 | g
[9 empty] | 0.97667 | g
[9 empty] | 0.91018 | g
[9 empty] | 0.9729 | g
[9 empty] | 0.9064 | g
[9 empty] | 0.99223 | g
[9 empty] | 0.92904 | g
[9 empty] | 0.93163 | g
[9 empty] | 0.99153 | g
[9 empty] | 0.97526 | g
[9 empty] | 0.96064 | g
[9 empty] | 0.93824 | g
[9 empty] | 0.93069 | g
[9 empty] | 0.91301 | g
[9 empty] | 0.96936 | g
[9 empty] | 0.95545 | g
[9 empty] | 0.96205 | g
[9 empty] | 0.97054 | g
[9 empty] | 0.93965 | g
[9 empty] | 0.97573 | g
[9 empty] | 0.90947 | g
[9 empty] | 0.93564 | g
[9 empty] | 0.54964 | g
[9 empty] | -0.02688 | g
[9 empty] | -0.07781 | g
[9 empty] | -0.01297 | g
[9 empty] | -0.02995 | g
[9 empty] | -0.08371 | g
[9 empty] | -0.02122 | g
[9 empty] | -0.08701 | g
[9 empty] | 2.004 | g
[9 empty] | 1.9349 | g
[9 empty] | 1.9819 | g
[9 empty] | 1.9614 | g
[9 empty] | 1.9866 | g
[9 empty] | 1.9963 | g
[9 empty] | 1.9668 | g
[9 empty] | 1.9458 | g
[9 empty] | 1.8996 | g
[9 empty] | 1.9034 | g
[9 empty] | 1.9001 | g
[9 empty] | 1.9672 | g
[9 empty] | 1.9658 | g
[9 empty] | 2.009 | g
[9 empty] | 1.9604 | g
[9 empty] | 1.9517 | g
[9 empty] | 1.9939 | g
[9 empty] | 1.9159 | g
[9 empty] | 1.9814 | g
[9 empty] | 1.9135 | g
[9 empty] | 1.9227 | g
[9 empty] | 1.9932 | g
[9 empty] | 1.9557 | g
[9 empty] | 1.9526 | g
[9 empty] | 2.0182 | g
[9 empty] | 1.9573 | g
[9 empty] | 2.0033 | g
[9 empty] | 2.0477 | g
[9 empty] | 1.9781 | g
[9 empty] | 2.05 | g
[9 empty] | 1.9927 | g
[9 empty] | 2.0017 | g
[9 empty] | 1.9743 | g
[9 empty] | 0.26362 | g
[9 empty] | -0.08842 | g
[9 empty] | -0.04952 | g
[9 empty] | 0.00872 | g
[9 empty] | -0.04716 | g
[9 empty] | -0.0099 | g
[9 empty] | 0.02122 | g
[9 empty] | 0.02358 | g
[9 empty] | 0.00896 | g
[9 empty] | -0.04056 | g
[9 empty] | 0.00542 | g
[9 empty] | -0.08111 | g
[9 empty] | -0.083 | g
[9 empty] | -0.09597 | g
[9 empty] | -0.02216 | g
[9 empty] | -0.02405 | g
[9 empty] | -0.04174 | g
[9 empty] | -0.02523 | g
[9 empty] | -0.01651 | g
[9 empty] | -0.00872 | g
[9 empty] | -0.03042 | g
[9 empty] | -0.02712 | g
[9 empty] | 4.6567 | g
[9 empty] | 4.9935 | g
[9 empty] | 4.9725 | g
[9 empty] | 4.9298 | g
[9 empty] | 4.9366 | g
[9 empty] | 4.9357 | g
[9 empty] | 4.9732 | g
[9 empty] | 4.9779 | g
[9 empty] | 4.9086 | g
[9 empty] | 4.9682 | g
[9 empty] | 4.9067 | g
[9 empty] | 4.9076 | g
[9 empty] | 4.9781 | g
[9 empty] | 4.9739 | g
[9 empty] | 4.9194 | g
[9 empty] | 4.9581 | g
[9 empty] | 4.977 | g
[9 empty] | 4.9383 | g
[9 empty] | 4.9557 | g
[9 empty] | 4.9949 | g
[9 empty] | 4.9902 | g
[9 empty] | 4.9246 | g
[9 empty] | 4.9517 | g
[9 empty] | 4.9152 | g
[9 empty] | 4.9411 | g
[9 empty] | 4.9326 | g
[9 empty] | 4.9541 | g
[9 empty] | 4.9671 | g
[9 empty] | 4.9777 | g
[9 empty] | 4.9939 | g
[9 empty] | 4.9008 | g
[9 empty] | 4.988 | g
[9 empty] | 4.9376 | g
[9 empty] | 4.9355 | g
[9 empty] | 4.9963 | g
[9 empty] | 4.9661 | g
[9 empty] | 4.9253 | g
[9 empty] | 4.9058 | g
[9 empty] | 1.5294 | g
[9 empty] | -0.06602 | g
[9 empty] | 0.06225 | g
[9 empty] | -0.04551 | g
[9 empty] | -0.01203 | g
[9 empty] | 0.02004 | g
[9 empty] | -0.00448 | g
[9 empty] | 3.6848 | g
[9 empty] | 9.9532 | g
[9 empty] | 9.9153 | g
[9 empty] | 9.9039 | g
[9 empty] | 9.941 | g
[9 empty] | 9.9546 | g
[9 empty] | 9.907 | g
[9 empty] | 9.9933 | g
[9 empty] | 9.9146 | g
[9 empty] | 9.8912 | g
[9 empty] | 9.9405 | g
[9 empty] | 9.9421 | g
[9 empty] | 9.9553 | g
[9 empty] | 9.9355 | g
[9 empty] | 9.9468 | g
[9 empty] | 9.9846 | g
[9 empty] | 9.9589 | g
[9 empty] | 9.9933 | g
[9 empty] | 9.9961 | g
[9 empty] | 9.948 | g
[9 empty] | 10.0164 | g
[9 empty] | 9.9792 | g
[9 empty] | 9.9549 | g
[9 empty] | 9.9212 | g
[9 empty] | 9.9766 | g
[9 empty] | 9.9888 | g
[9 empty] | 9.9907 | g
[9 empty] | 9.9169 | g
[9 empty] | 9.9749 | g
[9 empty] | 10.0584 | g
[9 empty] | 9.9535 | g
[9 empty] | 9.978 | g
[9 empty] | 14.8311 | g
[9 empty] | -0.08866 | g
[9 empty] | -0.04622 | g
[9 empty] | -0.05329 | g
[9 empty] | -0.05895 | g
[9 empty] | -0.04952 | g
[9 empty] | -0.04433 | g
[9 empty] | -0.04999 | g
[9 empty] | 0.0099 | g
[9 empty] | -0.02853 | g
[9 empty] | 9.7639 | g
[9 empty] | 19.9347 | g
[9 empty] | 19.9534 | g
[9 empty] | 19.9194 | g
[9 empty] | 19.9835 | g
[9 empty] | 19.9008 | g
[9 empty] | 19.9637 | g
[9 empty] | 20.005 | g
[9 empty] | 19.9286 | g
[9 empty] | 19.9762 | g
[9 empty] | 19.9385 | g
[9 empty] | 19.9281 | g
[9 empty] | 19.8777 | g
[9 empty] | 19.9788 | g
[9 empty] | 19.9786 | g
[9 empty] | 19.947 | g
[9 empty] | 19.9166 | g
[9 empty] | 19.9123 | g
[9 empty] | 19.9694 | g
[9 empty] | 19.9392 | g
[9 empty] | 19.9383 | g
[9 empty] | 19.9581 | g
[9 empty] | 20.0227 | g
[9 empty] | 19.9718 | g
[9 empty] | 19.9663 | g
[9 empty] | 19.9392 | g
[9 empty] | 19.9654 | g
[9 empty] | 20.0257 | g
[9 empty] | 19.9923 | g
[9 empty] | 20.0034 | g
[9 empty] | 19.9628 | g
[9 empty] | 19.9868 | g
[9 empty] | 19.9729 | g
[9 empty] | 19.9916 | g
[9 empty] | 20.0036 | g
[9 empty] | 20.3073 | g
[9 empty] | 2.4263 | g
[9 empty] | -0.07145 | g
[9 empty] | 0.00448 | g
[9 empty] | 0.02358 | g
[9 empty] | 0.02947 | g
[9 empty] | -0.06107 | g
[9 empty] | -0.08347 | g
[9 empty] | -0.03584 | g
[9 empty] | 15.3638 | g
[9 empty] | 50.0113 | g
[9 empty] | 50.0071 | g
[9 empty] | 49.9227 | g
[9 empty] | 49.9639 | g
[9 empty] | 49.9908 | g
[9 empty] | 49.9087 | g
[9 empty] | 49.9856 | g
[9 empty] | 49.9387 | g
[9 empty] | 49.9427 | g
[9 empty] | 49.8885 | g
[9 empty] | 49.9755 | g
[9 empty] | 49.9087 | g
[9 empty] | 49.9545 | g
[9 empty] | 49.9757 | g
[9 empty] | 49.9302 | g
[9 empty] | 49.988 | g
[9 empty] | 49.9679 | g
[9 empty] | 49.9057 | g
[9 empty] | 49.913 | g
[9 empty] | 49.9175 | g
[9 empty] | 49.971 | g
[9 empty] | 49.9531 | g
[9 empty] | 49.9061 | g
[9 empty] | 49.9696 | g
[9 empty] | 50.0196 | g
[9 empty] | 50.0158 | g
[9 empty] | 50.0203 | g
[9 empty] | 49.9887 | g
[9 empty] | 49.9776 | g
[9 empty] | 49.9594 | g
[9 empty] | 49.9587 | g
[9 empty] | 50.0163 | g
[9 empty] | 49.9646 | g
[9 empty] | 49.9488 | g
[9 empty] | 49.9326 | g
[9 empty] | 49.9191 | g
[9 empty] | 49.9087 | g
[9 empty] | 49.962 | g
[9 empty] | 50.5508 | g
[9 empty] | 12.5757 | g
[9 empty] | -0.03207 | g
[9 empty] | -0.00259 | g
[9 empty] | -0.03773 | g
[9 empty] | -0.02971 | g
[9 empty] | -0.05282 | g
[9 empty] | -0.09149 | g
[9 empty] | -0.03207 | g
[9 empty] | -0.02452 | g
[9 empty] | -0.08654 | g
[9 empty] | -0.01179 | g
[9 empty] | -0.02334 | g
[9 empty] | 23.8843 | g
[9 empty] | 99.8041 | g
[9 empty] | 99.9585 | g
[9 empty] | 99.992 | g
[9 empty] | 99.9502 | g
[9 empty] | 99.9557 | g
[9 empty] | 99.9639 | g
[9 empty] | 99.9903 | g
[9 empty] | 99.9437 | g
[9 empty] | 99.9958 | g
[9 empty] | 99.8602 | g
[9 empty] | 99.9875 | g
[9 empty] | 99.9715 | g
[9 empty] | 99.9069 | g
[9 empty] | 99.925 | g
[9 empty] | 99.9682 | g
[9 empty] | 99.9686 | g
[9 empty] | 99.9444 | g
[9 empty] | 99.9255 | g
[9 empty] | 99.9453 | g
[9 empty] | 99.9472 | g
[9 empty] | 99.9816 | g
[9 empty] | 99.9149 | g
[9 empty] | 99.9106 | g
[9 empty] | 99.9356 | g
[9 empty] | 99.904 | g
[9 empty] | 99.9377 | g
[9 empty] | 99.9729 | g
[9 empty] | 99.9564 | g
[9 empty] | 99.9543 | g
[9 empty] | 99.9267 | g
[9 empty] | 99.9625 | g
[9 empty] | 99.9215 | g
[9 empty] | 47.0351 | g
[9 empty] | 0.01698 | g
[9 empty] | -0.01085 | g
[9 empty] | 0.01108 | g
[9 empty] | -0.05494 | g
[9 empty] | -0.0349 | g
[9 empty] | -0.05518 | g
[9 empty] | -0.03608 | g
[9 empty] | -0.03372 | g
[9 empty] | -0.01108 | g
[9 empty] | -0.00424 | g
[9 empty] | 0.74441 | g
[9 empty] | 0.91866 | g
[9 empty] | 0.99718 | g
[9 empty] | 0.9538 | g
[9 empty] | 1.0071 | g
[9 empty] | 0.95285 | g
[9 empty] | 1.0094 | g
[9 empty] | 0.97408 | g
[9 empty] | 1.021 | g
[9 empty] | 0.98351 | g
[9 empty] | 0.99695 | g
[9 empty] | 0.938 | g
[9 empty] | 0.9637 | g
[9 empty] | 0.97974 | g
[9 empty] | 0.98822 | g
[9 empty] | 1.038 | g
[9 empty] | 1.0047 | g
[9 empty] | 0.96111 | g
[9 empty] | 0.96394 | g
[9 empty] | 0.95215 | g
[9 empty] | 0.90923 | g
[9 empty] | 0.94955 | g
[9 empty] | 1.0472 | g
[9 empty] | 0.94696 | g
[9 empty] | 0.99553 | g
[9 empty] | 0.98445 | g
[9 empty] | 1.0033 | g
[9 empty] | 1.0356 | g
[9 empty] | 1.0566 | g
[9 empty] | 0.98917 | g
[9 empty] | 1.0111 | g
[9 empty] | 0.99789 | g
[9 empty] | 0.99058 | g
[9 empty] | 0.96771 | g
[9 empty] | 0.93776 | g
[9 empty] | 1.4586 | g
[9 empty] | 0.01603 | g
[9 empty] | 0.01061 | g
[9 empty] | -0.01839 | g
[9 empty] | -0.05518 | g
[9 empty] | -0.10446 | g
[9 empty] | -0.04409 | g
[9 empty] | 0.42797 | g
[9 empty] | 2.0432 | g
[9 empty] | 2.0137 | g
[9 empty] | 1.938 | g
[9 empty] | 1.9672 | g
[9 empty] | 1.9955 | g
[9 empty] | 1.9762 | g
[9 empty] | 1.9746 | g
[9 empty] | 1.9597 | g
[9 empty] | 2.0333 | g
[9 empty] | 2.0701 | g
[9 empty] | 1.9866 | g
[9 empty] | 1.9609 | g
[9 empty] | 2.0184 | g
[9 empty] | 2.0842 | g
[9 empty] | 1.9937 | g
[9 empty] | 1.9929 | g
[9 empty] | 2.0293 | g
[9 empty] | 2.0189 | g
[9 empty] | 2.009 | g
[9 empty] | 1.9526 | g
[9 empty] | 1.9651 | g
[9 empty] | 2.009 | g
[9 empty] | 1.9713 | g
[9 empty] | 1.9856 | g
[9 empty] | 1.9977 | g
[9 empty] | 1.971 | g
[9 empty] | 2.001 | g
[9 empty] | 1.9731 | g
[9 empty] | 1.9562 | g
[9 empty] | 2.0533 | g
[9 empty] | 1.971 | g
[9 empty] | 1.9637 | g
[9 empty] | 1.9588 | g
[9 empty] | 7.2215 | g
[9 empty] | 0.01721 | g
[9 empty] | -0.02287 | g
[9 empty] | -0.0514 | g
[9 empty] | -0.06602 | g
[9 empty] | -0.01533 | g
[9 empty] | -0.029 | g
[9 empty] | 0.00047 | g
[9 empty] | 5.7391 | g
[9 empty] | 4.9781 | g
[9 empty] | 4.9715 | g
[9 empty] | 4.9588 | g
[9 empty] | 4.9392 | g
[9 empty] | 4.9555 | g
[9 empty] | 5.001 | g
[9 empty] | 4.9249 | g
[9 empty] | 4.9869 | g
[9 empty] | 4.9404 | g
[9 empty] | 5.0036 | g
[9 empty] | 4.9789 | g
[9 empty] | 4.963 | g
[9 empty] | 5.0003 | g
[9 empty] | 4.894 | g
[9 empty] | 4.9958 | g
[9 empty] | 4.9428 | g
[9 empty] | 5.002 | g
[9 empty] | 5.0281 | g
[9 empty] | 4.8925 | g
[9 empty] | 4.9411 | g
[9 empty] | 4.9793 | g
[9 empty] | 5.0156 | g
[9 empty] | 5.0079 | g
[9 empty] | 4.9208 | g
[9 empty] | 5.0145 | g
[9 empty] | 4.9692 | g
[9 empty] | 5.0352 | g
[9 empty] | 4.9906 | g
[9 empty] | 5.0229 | g
[9 empty] | 4.9614 | g
[9 empty] | 5.0156 | g
[9 empty] | 4.9805 | g
[9 empty] | 4.9963 | g
[9 empty] | 4.9876 | g
[9 empty] | 1.4143 | g
[9 empty] | -0.03042 | g
[9 empty] | -0.02146 | g
[9 empty] | -0.03301 | g
[9 empty] | 0.01415 | g
[9 empty] | -0.04504 | g
[9 empty] | -0.03891 | g
[9 empty] | -0.04504 | g
[9 empty] | 1.6617 | g
[9 empty] | 9.9431 | g
[9 empty] | 9.9549 | g
[9 empty] | 9.9766 | g
[9 empty] | 9.9412 | g
[9 empty] | 10.0155 | g
[9 empty] | 9.9537 | g
[9 empty] | 10.0013 | g
[9 empty] | 9.9806 | g
[9 empty] | 10.0362 | g
[9 empty] | 9.9693 | g
[9 empty] | 9.9846 | g
[9 empty] | 9.9957 | g
[9 empty] | 9.9617 | g
[9 empty] | 9.9358 | g
[9 empty] | 9.9822 | g
[9 empty] | 9.9747 | g
[9 empty] | 9.9563 | g
[9 empty] | 10.0336 | g
[9 empty] | 9.9716 | g
[9 empty] | 9.9902 | g
[9 empty] | 10.0178 | g
[9 empty] | 10.0237 | g
[9 empty] | 10.0006 | g
[9 empty] | 10.0122 | g
[9 empty] | 9.9629 | g
[9 empty] | 9.9563 | g
[9 empty] | 9.9627 | g
[9 empty] | 10.0006 | g
[9 empty] | 9.9634 | g
[9 empty] | 9.9551 | g
[9 empty] | 9.9877 | g
[9 empty] | 9.9961 | g
[9 empty] | 10.0056 | g
[9 empty] | 9.9518 | g
[9 empty] | 9.9339 | g
[9 empty] | 9.9785 | g
[9 empty] | 9.9789 | g
[9 empty] | 9.9799 | g
[9 empty] | 10.0233 | g
[9 empty] | 9.9822 | g
[9 empty] | 9.9829 | g
[9 empty] | 10.0499 | g
[9 empty] | 9.9542 | g
[9 empty] | 9.9487 | g
[9 empty] | 9.9657 | g
[9 empty] | 9.97 | g
[9 empty] | 9.9655 | g
[9 empty] | 7.5429 | g
[9 empty] | -0.04056 | g
[9 empty] | -0.00212 | g
[9 empty] | -0.03042 | g
[9 empty] | 0.01816 | g
[9 empty] | -0.054 | g
[9 empty] | 11.5807 | g
[9 empty] | 19.9498 | g
[9 empty] | 19.9656 | g
[9 empty] | 19.9741 | g
[9 empty] | 19.9255 | g
[9 empty] | 19.9607 | g
[9 empty] | 19.9689 | g
[9 empty] | 19.9305 | g
[9 empty] | 19.972 | g
[9 empty] | 19.9444 | g
[9 empty] | 19.9244 | g
[9 empty] | 19.9456 | g
[9 empty] | 19.9147 | g
[9 empty] | 20.0449 | g
[9 empty] | 19.9927 | g
[9 empty] | 19.9314 | g
[9 empty] | 19.9277 | g
[9 empty] | 19.9626 | g
[9 empty] | 20.0062 | g
[9 empty] | 19.9385 | g
[9 empty] | 20.0557 | g
[9 empty] | 19.9913 | g
[9 empty] | 19.9668 | g
[9 empty] | 19.9607 | g
[9 empty] | 20.0208 | g
[9 empty] | 19.9998 | g
[9 empty] | 19.959 | g
[9 empty] | 20.0354 | g
[9 empty] | 19.9937 | g
[9 empty] | 19.993 | g
[9 empty] | 19.9463 | g
[9 empty] | 19.972 | g
[9 empty] | 19.9534 | g
[9 empty] | 19.9331 | g
[9 empty] | 19.9524 | g
[9 empty] | 19.9946 | g
[9 empty] | 19.9703 | g
[9 empty] | 19.9845 | g
[9 empty] | 20.0559 | g
[9 empty] | 19.9842 | g
[9 empty] | 20.0239 | g
[9 empty] | 3.4702 | g
[9 empty] | 0.01344 | g
[9 empty] | -0.00684 | g
[9 empty] | -0.02853 | g
[9 empty] | -0.02216 | g
[9 empty] | 0.00307 | g
[9 empty] | 0.03985 | g
[9 empty] | -0.00684 | g
[9 empty] | -0.03183 | g
[9 empty] | -0.0283 | g
[9 empty] | -0.05494 | g
[9 empty] | 10.591 | g
[9 empty] | 50.0111 | g
[9 empty] | 49.9762 | g
[9 empty] | 49.9738 | g
[9 empty] | 49.9399 | g
[9 empty] | 49.9599 | g
[9 empty] | 49.9429 | g
[9 empty] | 49.9682 | g
[9 empty] | 49.9953 | g
[9 empty] | 49.9972 | g
[9 empty] | 49.9316 | g
[9 empty] | 49.9427 | g
[9 empty] | 50.046 | g
[9 empty] | 49.9703 | g
[9 empty] | 49.9399 | g
[9 empty] | 49.9974 | g
[9 empty] | 49.9271 | g
[9 empty] | 49.9538 | g
[9 empty] | 49.9762 | g
[9 empty] | 49.941 | g
[9 empty] | 49.9528 | g
[9 empty] | 50.0082 | g
[9 empty] | 49.9776 | g
[9 empty] | 49.9538 | g
[9 empty] | 49.9354 | g
[9 empty] | 49.9597 | g
[9 empty] | 49.9533 | g
[9 empty] | 49.946 | g
[9 empty] | 49.9667 | g
[9 empty] | 50.0028 | g
[9 empty] | 49.9734 | g
[9 empty] | 49.9392 | g
[9 empty] | 49.9906 | g
[9 empty] | 50.0434 | g
[9 empty] | 49.9875 | g
[9 empty] | 49.922 | g
[9 empty] | 49.9512 | g
[9 empty] | 49.9863 | g
[9 empty] | 49.967 | g
[9 empty] | 49.9286 | g
[9 empty] | 49.9599 | g
[9 empty] | 49.9057 | g
[9 empty] | 49.983 | g
[9 empty] | 49.9125 | g
[9 empty] | 49.9936 | g
[9 empty] | 33.3143 | g
[9 empty] | -0.00141 | g
[9 empty] | 0.00684 | g
[9 empty] | -0.03136 | g
[9 empty] | 0.00189 | g
[9 empty] | -0.01745 | g
[9 empty] | -0.03537 | g
[9 empty] | -0.04952 | g
[9 empty] | -0.02287 | g
[9 empty] | -0.00896 | g
[9 empty] | -0.01886 | g
[9 empty] | -0.03938 | g
[9 empty] | 58.6613 | g
[9 empty] | 99.9194 | g
[9 empty] | 100.0123 | g
[9 empty] | 99.9387 | g
[9 empty] | 99.9109 | g
[9 empty] | 99.9658 | g
[9 empty] | 99.9003 | g
[9 empty] | 99.9583 | g
[9 empty] | 99.897 | g
[9 empty] | 99.883 | g
[9 empty] | 99.9328 | g
[9 empty] | 99.9399 | g
[9 empty] | 99.917 | g
[9 empty] | 99.951 | g
[9 empty] | 99.9288 | g
[9 empty] | 99.9356 | g
[9 empty] | 99.9517 | g
[9 empty] | 99.9863 | g
[9 empty] | 99.9712 | g
[9 empty] | 99.9533 | g
[9 empty] | 99.9743 | g
[9 empty] | 99.9243 | g
[9 empty] | 99.9547 | g
[9 empty] | 99.9121 | g
[9 empty] | 100.0073 | g
[9 empty] | 100.0193 | g
[9 empty] | 100.0182 | g
[9 empty] | 99.9361 | g
[9 empty] | 99.971 | g
[9 empty] | 100.0264 | g
[9 empty] | 99.9552 | g
[9 empty] | 99.9359 | g
[9 empty] | 99.9474 | g
[9 empty] | 99.9727 | g
[9 empty] | 99.963 | g
[9 empty] | 99.9587 | g
[9 empty] | 99.9691 | g
[9 empty] | 99.9823 | g
[9 empty] | 99.9802 | g
[9 empty] | 67.2415 | g
[9 empty] | 0.0191 | g
[9 empty] | 0.01226 | g
[9 empty] | 0.05966 | g
[9 empty] | -0.04339 | g
[9 empty] | 0.01273 | g
[9 empty] | -0.01108 | g
[9 empty] | 0.01226 | g
[9 empty] | -0.0158 | g
[9 empty] | -0.07758 | g
[9 empty] | -0.02688 | g
[9 empty] | -0.01344 | g
[9 empty] | -0.01721 | g
[9 empty] | -0.029 | g
[9 empty] | -0.05706 | g
[9 empty] | 0 | g
[9 empty] | 0.0316 | g
[9 empty] | -0.0191 | g
[9 empty] | -0.03726 | g
[9 empty] | 0.02216 | g
[9 empty] | 0.03372 | g
[9 empty] | -0.01745 | g
[9 empty] | -0.01816 | g
[9 empty] | 0.03867 | g
[9 empty] | 0.04787 | g
[9 empty] | 0.00542 | g
[9 empty] | 0.01745 | g
[9 empty] | 0.01297 | g
[9 empty] | 0.0033 | g
[9 empty] | -0.01014 | g
[9 empty] | -0.01651 | g
[9 empty] | -0.01415 | g
[9 empty] | 0.02947 | g
[9 empty] | -0.01438 | g
[9 empty] | 0.02735 | g
[9 empty] | -0.01344 | g
[9 empty] | -0.05164 | g
[9 empty] | 0.0382 | g
[9 empty] | 0.02311 | g
[9 empty] | -0.00896 | g
[9 empty] | -0.01038 | g
[9 empty] | -0.01627 | g
[9 empty] | -0.03136 | g
[9 empty] | -0.00189 | g
[9 empty] | -0.05188 | g
[9 empty] | -0.02641 | g
[9 empty] | -0.02264 | g
[9 empty] | -0.00071 | g
[9 empty] | -0.02688 | g
[9 empty] | -0.06036 | g
[9 empty] | 0.00731 | g
[9 empty] | -0.02499 | g
[9 empty] | -0.01792 | g
[9 empty] | -0.05942 | g
[9 empty] | 0.01627 | g
[9 empty] | -0.0481 | g
[9 empty] | -0.04197 | g
[9 empty] | -0.0573 | g
[9 empty] | 0.08654 | g
[9 empty] | -0.00872 | g
[9 empty] | -0.02358 | g
[9 empty] | 0.01533 | g
[9 empty] | -0.02004 | g
[9 empty] | 0.02311 | g

## 200624 - 5000g Standards PL (2)
Mass Balance Test
Device Under Test: | OS_Nano_Balance (3D printed balance), Using TAL220 5kg load cell
Lab Balance: | N/A Using standard masses
DUT Calibration: | Calibrated with 100 g standard mass
| 424.094 | g/div
PuTTY
[8 empty] | 1G | 0.3096 | g
|  |  | Nominal | Average | Std | Count | Addresses | 1G | 0.3445 | g
|  |  | 1 | 1.0061 | 0.034448 | 37 | S86:S122 | 1G | 0.27777 | g
|  |  | 2 | 2.0626 | 0.044518 | 42 | S139:S180 | 1G | 0.2997 | g
|  |  | 5 | 5.1094 | 0.040104 | 48 | S191:S238 | 1G | 0.26787 | g
|  |  | 10 | 10.1323 | 0.039517 | 34 | S253:S286 | 1G | 0.23273 | g
|  |  | 20 | 19.9692 | 0.037422 | 35 | S318:S352 | 1G | 0.30866 | g
|  |  | 50 | 50.0191 | 0.033414 | 37 | S365:S401 | 1G | 0.28791 | g
|  |  | 100 | 100.0451 | 0.034799 | 35 | S419:S453 | 1G | 0.29758 | g
|  |  | 1 | 1.0498 | 0.039107 | 40 | S513:S552 | 1G | 0.3254 | g
|  |  | 2 | 2.099 | 0.034504 | 43 | S566:S608 | 1G | 0.32941 | g
|  |  | 5 | 5.1156 | 0.036657 | 35 | S618:S652 | 1G | 0.26621 | g
|  |  | 10 | 10.1834 | 0.033048 | 32 | S662:S693 | 1G | 0.35016 | g
|  |  | 20 | 20.2395 | 0.037647 | 35 | S700:S734 | 1G | 0.31927 | g
|  |  | 50 | 50.2969 | 0.032689 | 39 | S760:S798 | 1G | 0.28083 | g
|  |  | 100 | 100.0714 | 0.033514 | 42 | S1196:S1237 | 1G | 0.31526 | g
[8 empty] | 1G | 0.34945 | g
|  |  |  | Avg | 0.036528 |  |  | 1G | 0.25938 | g
|  |  |  | Max | 0.044518 |  |  | 1G | 0.2516 | g
[8 empty] | 1G | 0.38175 | g
[8 empty] | 1G | 0.30559 | g
[8 empty] | 1G | 0.34238 | g
[8 empty] | 1G | 0.3985 | g
[8 empty] | 1G | 0.36454 | g
[8 empty] | 1G | 0.37303 | g
[8 empty] | 1G | 0.29121 | g
[8 empty] | 1G | 0.32823 | g
[8 empty] | 1G | 0.33082 | g
[8 empty] | 1G | 0.34686 | g
[8 empty] | 1G | 0.37445 | g
[8 empty] | 1G | 0.37445 | g
[8 empty] | 1G | 0.27187 | g
[8 empty] | 1G | 0.35723 | g
[8 empty] | 1G | 0.34804 | g
[8 empty] | 1G | 0.38718 | g
^J
[8 empty] | 1G | 0.37987 | g
z
[8 empty] | 1G | 0.29357 | g
[8 empty] | 1G | 0.36171 | g
[8 empty] | Z1G | -0.05895 | g
^Jr
[8 empty] | 1G | -0.02452 | g
[8 empty] | 1G | -0.08159 | g
[8 empty] | 1G | -0.0514 | g
[8 empty] | 1G | -0.03042 | g
[8 empty] | 1G | -0.07215 | g
[8 empty] | 1G | -0.02735 | g
[8 empty] | 1G | -0.04692 | g
[8 empty] | 1G | -0.00896 | g
[8 empty] | 1G | -0.00189 | g
[8 empty] | 1G | -0.05188 | g
[8 empty] | 1G | 0.01556 | g
^J
[8 empty] | 1G | 0.01344 | g
[8 empty] | 1G | 0.02617 | g
[8 empty] | 1G | 0.05447 | g
[8 empty] | 1G | 0.00448 | g
[8 empty] | 1G | -0.00236 | g
[8 empty] | 1G | 0.02594 | g
[8 empty] | 1G | -0.03891 | g
[8 empty] | 1G | 0.03395 | g
r
[8 empty] | 1G | 0.01014 | g
[8 empty] | 1G | 0.01344 | g
[8 empty] | 1G | 0.05942 | g
[8 empty] | 1G | 0.04763 | g
[8 empty] | 1G | 0.04504 | g
[8 empty] | 1G | 0.03325 | g
[8 empty] | 1G | 1.0651 | g
[8 empty] | 1G | 1.0451 | g
[8 empty] | 1G | 1.0288 | g
[8 empty] | 1G | 0.96347 | g
[8 empty] | 1G | 0.94319 | g
[8 empty] | 1G | 0.95993 | g
[8 empty] | 1G | 1.0825 | g
[8 empty] | 1G | 0.96842 | g
[8 empty] | 1G | 1.0481 | g
[8 empty] | 1G | 1.0054 | g
[8 empty] | 1G | 0.95733 | g
[8 empty] | 1G | 1.0507 | g
[8 empty] | 1G | 1 | g
[8 empty] | 1G | 0.97974 | g
[8 empty] | 1G | 1.0153 | g
[8 empty] | 1G | 1.0142 | g
[8 empty] | 1G | 0.9663 | g
[8 empty] | 1G | 1.0156 | g
[8 empty] | 1G | 0.98115 | g
[8 empty] | 1G | 1.0163 | g
[8 empty] | 1G | 0.95804 | g
[8 empty] | 1G | 1.0455 | g
[8 empty] | 1G | 1.0071 | g
[8 empty] | 1G | 1.041 | g
[8 empty] | 1G | 1.008 | g
[8 empty] | 1G | 1.07 | g
[8 empty] | 1G | 0.97502 | g
[8 empty] | 1G | 0.9927 | g
[8 empty] | 1G | 0.98539 | g
[8 empty] | 1G | 1.0347 | g
[8 empty] | 1G | 1.0224 | g
[8 empty] | 1G | 0.96535 | g
[8 empty] | 1G | 1.0314 | g
[8 empty] | 1G | 0.9696 | g
[8 empty] | 1G | 1.0094 | g
[8 empty] | 1G | 1.0453 | g
[8 empty] | 1G | 1.0219 | g
[8 empty] | 1G | 1.0014 | g
[8 empty] | 1G | 0.99058 | g
[8 empty] | 1G | 1.3945 | g
[8 empty] | 1G | 0.00307 | g
[8 empty] | 1G | 0.02193 | g
[8 empty] | 1G | -0.00542 | g
[8 empty] | 1G | -0.03513 | g
[8 empty] | 1G | -0.06862 | g
[8 empty] | 1G | 0.0448 | g
[8 empty] | 1G | -0.06084 | g
[8 empty] | 1G | -0.02594 | g
[8 empty] | 1G | 0.00825 | g
[8 empty] | 1G | 0.02947 | g
[8 empty] | 1G | -0.03089 | g
[8 empty] | 1G | 1.4228 | g
[8 empty] | 1G | 1.9456 | g
[8 empty] | 1G | 1.9868 | g
[8 empty] | 1G | 1.9944 | g
[8 empty] | 1G | 2.0083 | g
[8 empty] | 1G | 2.0135 | g
[8 empty] | 1G | 1.9838 | g
[8 empty] | 1G | 2.0047 | g
[8 empty] | 1G | 1.9519 | g
[8 empty] | 1G | 1.996 | g
[8 empty] | 1G | 2.0236 | g
[8 empty] | 1G | 2.0111 | g
[8 empty] | 1G | 2.0899 | g
[8 empty] | 1G | 2.0569 | g
[8 empty] | 1G | 2.079 | g
[8 empty] | 1G | 2.0097 | g
[8 empty] | 1G | 2.0811 | g
[8 empty] | 1G | 2.0375 | g
[8 empty] | 1G | 2.0472 | g
[8 empty] | 1G | 2.0868 | g
[8 empty] | 1G | 2.08 | g
[8 empty] | 1G | 2.1104 | g
[8 empty] | 1G | 2.0198 | g
[8 empty] | 1G | 2.1052 | g
[8 empty] | 1G | 2.0762 | g
[8 empty] | 1G | 2.1276 | g
[8 empty] | 1G | 2.0833 | g
[8 empty] | 1G | 2.1111 | g
[8 empty] | 1G | 2.0439 | g
[8 empty] | 1G | 2.116 | g
[8 empty] | 1G | 2.0951 | g
[8 empty] | 1G | 2.1267 | g
[8 empty] | 1G | 2.1427 | g
[8 empty] | 1G | 2.0533 | g
[8 empty] | 1G | 2.0908 | g
[8 empty] | 1G | 2.0785 | g
[8 empty] | 1G | 2.1005 | g
[8 empty] | 1G | 2.0243 | g
[8 empty] | 1G | 2.072 | g
[8 empty] | 1G | 2.0491 | g
[8 empty] | 1G | 2.0795 | g
[8 empty] | 1G | 2.0965 | g
[8 empty] | 1G | 2.0951 | g
[8 empty] | 1G | 2.0698 | g
[8 empty] | 1G | 2.1087 | g
[8 empty] | 1G | 1.9081 | g
[8 empty] | 1G | 0.08772 | g
[8 empty] | 1G | 0.08276 | g
[8 empty] | 1G | 0.11389 | g
[8 empty] | 1G | 0.10799 | g
[8 empty] | 1G | 0.03961 | g
[8 empty] | 1G | 0.06272 | g
[8 empty] | 1G | 0.13488 | g
[8 empty] | 1G | 0.5487 | g
[8 empty] | 1G | 5.1168 | g
[8 empty] | 1G | 5.1005 | g
[8 empty] | 1G | 5.1107 | g
[8 empty] | 1G | 5.1166 | g
[8 empty] | 1G | 4.9838 | g
[8 empty] | 1G | 5.0463 | g
[8 empty] | 1G | 5.0923 | g
[8 empty] | 1G | 5.0647 | g
[8 empty] | 1G | 5.1479 | g
[8 empty] | 1G | 5.1109 | g
[8 empty] | 1G | 5.0633 | g
[8 empty] | 1G | 5.1088 | g
[8 empty] | 1G | 5.0946 | g
[8 empty] | 1G | 5.1314 | g
[8 empty] | 1G | 5.0814 | g
[8 empty] | 1G | 5.0923 | g
[8 empty] | 1G | 5.0845 | g
[8 empty] | 1G | 5.1387 | g
[8 empty] | 1G | 5.0185 | g
[8 empty] | 1G | 5.1029 | g
[8 empty] | 1G | 5.1177 | g
[8 empty] | 1G | 5.0715 | g
[8 empty] | 1G | 5.1323 | g
[8 empty] | 1G | 5.0668 | g
[8 empty] | 1G | 5.1354 | g
[8 empty] | 1G | 5.1159 | g
[8 empty] | 1G | 5.1173 | g
[8 empty] | 1G | 5.0642 | g
[8 empty] | 1G | 5.0288 | g
[8 empty] | 1G | 5.1227 | g
[8 empty] | 1G | 5.1048 | g
[8 empty] | 1G | 5.1243 | g
[8 empty] | 1G | 5.1581 | g
[8 empty] | 1G | 5.1434 | g
[8 empty] | 1G | 5.0953 | g
[8 empty] | 1G | 5.1437 | g
[8 empty] | 1G | 5.1128 | g
[8 empty] | 1G | 5.1359 | g
[8 empty] | 1G | 5.1163 | g
[8 empty] | 1G | 5.1189 | g
[8 empty] | 1G | 5.1267 | g
[8 empty] | 1G | 5.1786 | g
[8 empty] | 1G | 5.1107 | g
[8 empty] | 1G | 5.1118 | g
[8 empty] | 1G | 5.1611 | g
[8 empty] | 1G | 5.1458 | g
[8 empty] | 1G | 5.1925 | g
[8 empty] | 1G | 5.1715 | g
[8 empty] | 1G | 5.1349 | g
[8 empty] | 1G | 5.0555 | g
[8 empty] | 1G | 0.10564 | g
[8 empty] | 1G | 0.19595 | g
[8 empty] | 1G | 0.16624 | g
[8 empty] | 1G | 0.10611 | g
[8 empty] | 1G | 0.15916 | g
[8 empty] | 1G | 0.09927 | g
[8 empty] | 1G | 0.14266 | g
[8 empty] | 1G | 0.18722 | g
[8 empty] | 1G | 0.12497 | g
[8 empty] | 1G | 0.13841 | g
[8 empty] | 1G | 0.99601 | g
[8 empty] | 1G | 6.6634 | g
[8 empty] | 1G | 10.1253 | g
[8 empty] | 1G | 10.0949 | g
[8 empty] | 1G | 10.0591 | g
[8 empty] | 1G | 10.1088 | g
[8 empty] | 1G | 10.1485 | g
[8 empty] | 1G | 10.1121 | g
[8 empty] | 1G | 10.1449 | g
[8 empty] | 1G | 10.1404 | g
[8 empty] | 1G | 10.1128 | g
[8 empty] | 1G | 10.1261 | g
[8 empty] | 1G | 10.0902 | g
[8 empty] | 1G | 10.1013 | g
[8 empty] | 1G | 10.1294 | g
[8 empty] | 1G | 10.0921 | g
[8 empty] | 1G | 10.1371 | g
[8 empty] | 1G | 10.1279 | g
[8 empty] | 1G | 10.0961 | g
[8 empty] | 1G | 10.0994 | g
[8 empty] | 1G | 10.0869 | g
[8 empty] | 1G | 10.0683 | g
[8 empty] | 1G | 10.1041 | g
[8 empty] | 1G | 10.1864 | g
[8 empty] | 1G | 10.1289 | g
[8 empty] | 1G | 10.1591 | g
[8 empty] | 1G | 10.1996 | g
[8 empty] | 1G | 10.1515 | g
[8 empty] | 1G | 10.1268 | g
[8 empty] | 1G | 10.1492 | g
[8 empty] | 1G | 10.1812 | g
[8 empty] | 1G | 10.1315 | g
[8 empty] | 1G | 10.2322 | g
[8 empty] | 1G | 10.1508 | g
[8 empty] | 1G | 10.1324 | g
[8 empty] | 1G | 10.222 | g
[8 empty] | 1G | 10.1654 | g
[8 empty] | 1G | 15.1714 | g
[8 empty] | 1G | 0.23886 | g
[8 empty] | 1G | 0.15893 | g
[8 empty] | 1G | 0.14502 | g
[8 empty] | 1G | 0.23273 | g
[8 empty] | 1G | 0.18204 | g
[8 empty] | 1G | 0.15185 | g
[8 empty] | 1G | 0.14973 | g
[8 empty] | 1G | 0.16883 | g
[8 empty] | 1G | 0.19736 | g
[8 empty] | 1G | 0.16341 | g
^Jz
[8 empty] | Z1G | -0.11082 | g
^Jr
[8 empty] | 1G | -0.10988 | g
[8 empty] | 1G | -0.09951 | g
[8 empty] | 1G | -0.01698 | g
[8 empty] | 1G | -0.04905 | g
[8 empty] | 1G | -0.08182 | g
[8 empty] | 1G | -0.04834 | g
[8 empty] | 1G | -0.07734 | g
[8 empty] | 1G | -0.02641 | g
[8 empty] | 1G | -0.11271 | g
[8 empty] | 1G | -0.06626 | g
[8 empty] | 1G | -0.06272 | g
[8 empty] | 1G | -0.03278 | g
[8 empty] | 1G | 0.41312 | g
[8 empty] | 1G | 18.2481 | g
[8 empty] | 1G | 19.9036 | g
[8 empty] | 1G | 19.9942 | g
[8 empty] | 1G | 19.9517 | g
[8 empty] | 1G | 19.942 | g
[8 empty] | 1G | 19.9043 | g
[8 empty] | 1G | 19.8911 | g
[8 empty] | 1G | 19.947 | g
[8 empty] | 1G | 19.9411 | g
[8 empty] | 1G | 19.9524 | g
[8 empty] | 1G | 19.9932 | g
[8 empty] | 1G | 19.9812 | g
[8 empty] | 1G | 19.9239 | g
[8 empty] | 1G | 19.9149 | g
[8 empty] | 1G | 19.9173 | g
[8 empty] | 1G | 20.0069 | g
[8 empty] | 1G | 19.9687 | g
[8 empty] | 1G | 19.9442 | g
[8 empty] | 1G | 20.0118 | g
[8 empty] | 1G | 19.9798 | g
[8 empty] | 1G | 19.9362 | g
[8 empty] | 1G | 19.9602 | g
[8 empty] | 1G | 19.9967 | g
[8 empty] | 1G | 20.0293 | g
[8 empty] | 1G | 19.9953 | g
[8 empty] | 1G | 19.9154 | g
[8 empty] | 1G | 20.0359 | g
[8 empty] | 1G | 19.9609 | g
[8 empty] | 1G | 20.0095 | g
[8 empty] | 1G | 19.9305 | g
[8 empty] | 1G | 20.0008 | g
[8 empty] | 1G | 19.9831 | g
[8 empty] | 1G | 19.9918 | g
[8 empty] | 1G | 20.0024 | g
[8 empty] | 1G | 19.9939 | g
[8 empty] | 1G | 19.9986 | g
[8 empty] | 1G | 20.0144 | g
[8 empty] | 1G | 19.8885 | g
[8 empty] | 1G | 0.50649 | g
[8 empty] | 1G | 0.01297 | g
[8 empty] | 1G | 0.05588 | g
[8 empty] | 1G | 0.02547 | g
[8 empty] | 1G | -0.01698 | g
[8 empty] | 1G | 0.11436 | g
[8 empty] | 1G | -0.04952 | g
[8 empty] | 1G | 0.00802 | g
[8 empty] | 1G | 0.07027 | g
[8 empty] | 1G | 5.2769 | g
[8 empty] | 1G | 51.515 | g
[8 empty] | 1G | 49.9948 | g
[8 empty] | 1G | 49.938 | g
[8 empty] | 1G | 49.9835 | g
[8 empty] | 1G | 49.9571 | g
[8 empty] | 1G | 50.033 | g
[8 empty] | 1G | 50.0752 | g
[8 empty] | 1G | 50.0106 | g
[8 empty] | 1G | 49.9601 | g
[8 empty] | 1G | 50.0417 | g
[8 empty] | 1G | 50.0177 | g
[8 empty] | 1G | 50.0066 | g
[8 empty] | 1G | 50.0318 | g
[8 empty] | 1G | 50.0307 | g
[8 empty] | 1G | 49.987 | g
[8 empty] | 1G | 49.987 | g
[8 empty] | 1G | 50.0224 | g
[8 empty] | 1G | 49.9903 | g
[8 empty] | 1G | 50.0453 | g
[8 empty] | 1G | 50.0391 | g
[8 empty] | 1G | 49.9943 | g
[8 empty] | 1G | 50.0691 | g
[8 empty] | 1G | 50.0455 | g
[8 empty] | 1G | 50.0729 | g
[8 empty] | 1G | 50.0038 | g
[8 empty] | 1G | 50.0212 | g
[8 empty] | 1G | 50.0139 | g
[8 empty] | 1G | 50.0648 | g
[8 empty] | 1G | 50.0571 | g
[8 empty] | 1G | 50.062 | g
[8 empty] | 1G | 50.0172 | g
[8 empty] | 1G | 50.0021 | g
[8 empty] | 1G | 50.0422 | g
[8 empty] | 1G | 50.0198 | g
[8 empty] | 1G | 49.9939 | g
[8 empty] | 1G | 49.9833 | g
[8 empty] | 1G | 50.0639 | g
[8 empty] | 1G | 50.0288 | g
[8 empty] | 1G | 14.7682 | g
[8 empty] | 1G | 0.09479 | g
[8 empty] | 1G | 0.00212 | g
[8 empty] | 1G | 0.01792 | g
[8 empty] | 1G | 0.06508 | g
[8 empty] | 1G | 0.0382 | g
[8 empty] | 1G | 0.02382 | g
[8 empty] | 1G | 0.07711 | g
[8 empty] | 1G | 0.02759 | g
[8 empty] | 1G | 0.112 | g
[8 empty] | 1G | 0.04763 | g
[8 empty] | 1G | -0.02193 | g
[8 empty] | 1G | 0.08394 | g
[8 empty] | 1G | 0.05636 | g
[8 empty] | 1G | 0.04857 | g
[8 empty] | 1G | 16.8944 | g
[8 empty] | 1G | 100.1775 | g
[8 empty] | 1G | 100.042 | g
[8 empty] | 1G | 100.0035 | g
[8 empty] | 1G | 100.0575 | g
[8 empty] | 1G | 100.0637 | g
[8 empty] | 1G | 99.9882 | g
[8 empty] | 1G | 100.0731 | g
[8 empty] | 1G | 100.0792 | g
[8 empty] | 1G | 100.0738 | g
[8 empty] | 1G | 99.9788 | g
[8 empty] | 1G | 99.991 | g
[8 empty] | 1G | 99.9908 | g
[8 empty] | 1G | 100 | g
[8 empty] | 1G | 100.0429 | g
[8 empty] | 1G | 100.0514 | g
[8 empty] | 1G | 100.0311 | g
[8 empty] | 1G | 100.0047 | g
[8 empty] | 1G | 100.0156 | g
[8 empty] | 1G | 100.0396 | g
[8 empty] | 1G | 100.1073 | g
[8 empty] | 1G | 99.9915 | g
[8 empty] | 1G | 100.0318 | g
[8 empty] | 1G | 100.0816 | g
[8 empty] | 1G | 100.0271 | g
[8 empty] | 1G | 100.0512 | g
[8 empty] | 1G | 100.0696 | g
[8 empty] | 1G | 100.0344 | g
[8 empty] | 1G | 100.1016 | g
[8 empty] | 1G | 100.0776 | g
[8 empty] | 1G | 100.054 | g
[8 empty] | 1G | 100.0823 | g
[8 empty] | 1G | 100.0561 | g
[8 empty] | 1G | 100.0521 | g
[8 empty] | 1G | 100.0599 | g
[8 empty] | 1G | 100.0764 | g
[8 empty] | 1G | 100.0967 | g
[8 empty] | 1G | 32.5742 | g
[8 empty] | 1G | 0.08536 | g
[8 empty] | 1G | 0.06201 | g
[8 empty] | 1G | 0.05636 | g
[8 empty] | 1G | 0.08913 | g
[8 empty] | 1G | 0.0929 | g
[8 empty] | 1G | 0.12096 | g
[8 empty] | 1G | 0.10186 | g
[8 empty] | 1G | 0.1719 | g
[8 empty] | 1G | 1.042 | g
[8 empty] | 1G | 1.1863 | g
[8 empty] | 1G | 1.1672 | g
[8 empty] | 1G | 1.0969 | g
[8 empty] | 1G | 1.1059 | g
[8 empty] | 1G | 1.2089 | g
[8 empty] | 1G | 1.1514 | g
[8 empty] | 1G | 9.3958 | g
[8 empty] | 1G | 0.09951 | g
[8 empty] | 1G | 0.15232 | g
[8 empty] | 1G | 0.14973 | g
[8 empty] | 1G | 0.09621 | g
[8 empty] | 1G | 0.15303 | g
[8 empty] | 1G | 0.05918 | g
[8 empty] | 1G | 0.17614 | g
[8 empty] | 1G | 0.13488 | g
[8 empty] | 1G | 0.16364 | g
[8 empty] | 1G | 0.15893 | g
[8 empty] | 1G | 0.14478 | g
[8 empty] | 1G | 0.09503 | g
^J
[8 empty] | 1G | 0.09738 | g
z
[8 empty] | 1G | 0.13605 | g
[8 empty] | Z1G | 0.01886 | g
^Jr
[8 empty] | 1G | 0.11035 | g
[8 empty] | 1G | 0.1495 | g
[8 empty] | 1G | 0.00354 | g
[8 empty] | 1G | 0.10116 | g
[8 empty] | 1G | 0.06319 | g
[8 empty] | 1G | 0.06154 | g
^J
[8 empty] | 1G | 0.00283 | g
z
[8 empty] | Z1G | 0.01179 | g
[8 empty] | j | ^Jr
[8 empty] | 1G | -0.00377 | g
[8 empty] | 1G | -0.0132 | g
[8 empty] | 1G | 0.04009 | g
[8 empty] | 1G | -0.03254 | g
[8 empty] | 1G | -0.00849 | g
[8 empty] | 1G | 0.00047 | g
[8 empty] | 1G | 0.03678 | g
[8 empty] | 1G | 1.0375 | g
[8 empty] | 1G | 1.0241 | g
[8 empty] | 1G | 1.058 | g
[8 empty] | 1G | 0.98728 | g
[8 empty] | 1G | 1.0175 | g
[8 empty] | 1G | 0.92645 | g
[8 empty] | 1G | 1.0337 | g
[8 empty] | 1G | 1.0352 | g
[8 empty] | 1G | 1.0505 | g
[8 empty] | 1G | 1.0486 | g
[8 empty] | 1G | 1.0389 | g
[8 empty] | 1G | 1.025 | g
[8 empty] | 1G | 1.0146 | g
[8 empty] | 1G | 1.0894 | g
[8 empty] | 1G | 1.0816 | g
[8 empty] | 1G | 1.0238 | g
[8 empty] | 1G | 1.0785 | g
[8 empty] | 1G | 1.0745 | g
[8 empty] | 1G | 1.0335 | g
[8 empty] | 1G | 1.0847 | g
[8 empty] | 1G | 1.0458 | g
[8 empty] | 1G | 1.0656 | g
[8 empty] | 1G | 1.0908 | g
[8 empty] | 1G | 1.0568 | g
[8 empty] | 1G | 1.0083 | g
[8 empty] | 1G | 1.0344 | g
[8 empty] | 1G | 1.0715 | g
[8 empty] | 1G | 1.1259 | g
[8 empty] | 1G | 1.0752 | g
[8 empty] | 1G | 1.1269 | g
[8 empty] | 1G | 1.0877 | g
[8 empty] | 1G | 1.0224 | g
[8 empty] | 1G | 1.0641 | g
[8 empty] | 1G | 1.0495 | g
[8 empty] | 1G | 1.0382 | g
[8 empty] | 1G | 0.99035 | g
[8 empty] | 1G | 1.0609 | g
[8 empty] | 1G | 1.0505 | g
[8 empty] | 1G | 1.0255 | g
[8 empty] | 1G | 1.1389 | g
[8 empty] | 1G | 1.0382 | g
[8 empty] | 1G | 0.58643 | g
[8 empty] | 1G | 0.00755 | g
[8 empty] | 1G | 0.08371 | g
[8 empty] | 1G | 0.02476 | g
[8 empty] | 1G | 0.0547 | g
[8 empty] | 1G | 0.0698 | g
[8 empty] | 1G | 0.0705 | g
[8 empty] | 1G | 0.0797 | g
[8 empty] | 1G | 0.06508 | g
[8 empty] | 1G | 0.00307 | g
[8 empty] | 1G | 0.03372 | g
[8 empty] | 1G | 0.3063 | g
[8 empty] | 1G | 2.1241 | g
[8 empty] | 1G | 2.1035 | g
[8 empty] | 1G | 2.0392 | g
[8 empty] | 1G | 2.0545 | g
[8 empty] | 1G | 2.0903 | g
[8 empty] | 1G | 2.0708 | g
[8 empty] | 1G | 2.0844 | g
[8 empty] | 1G | 2.0639 | g
[8 empty] | 1G | 2.129 | g
[8 empty] | 1G | 2.0366 | g
[8 empty] | 1G | 2.0793 | g
[8 empty] | 1G | 2.1226 | g
[8 empty] | 1G | 2.1068 | g
[8 empty] | 1G | 2.0097 | g
[8 empty] | 1G | 2.1441 | g
[8 empty] | 1G | 2.0309 | g
[8 empty] | 1G | 2.0755 | g
[8 empty] | 1G | 2.0894 | g
[8 empty] | 1G | 2.1165 | g
[8 empty] | 1G | 2.1068 | g
[8 empty] | 1G | 2.1326 | g
[8 empty] | 1G | 2.0951 | g
[8 empty] | 1G | 2.1137 | g
[8 empty] | 1G | 2.1172 | g
[8 empty] | 1G | 2.0748 | g
[8 empty] | 1G | 2.0884 | g
[8 empty] | 1G | 2.088 | g
[8 empty] | 1G | 2.1007 | g
[8 empty] | 1G | 2.112 | g
[8 empty] | 1G | 2.0649 | g
[8 empty] | 1G | 2.101 | g
[8 empty] | 1G | 2.1054 | g
[8 empty] | 1G | 2.1132 | g
[8 empty] | 1G | 2.109 | g
[8 empty] | 1G | 2.1788 | g
[8 empty] | 1G | 2.1835 | g
[8 empty] | 1G | 2.1097 | g
[8 empty] | 1G | 2.0807 | g
[8 empty] | 1G | 2.1373 | g
[8 empty] | 1G | 2.1191 | g
[8 empty] | 1G | 2.125 | g
[8 empty] | 1G | 2.1198 | g
[8 empty] | 1G | 2.1132 | g
[8 empty] | 1G | 2.1217 | g
[8 empty] | 1G | 0.5652 | g
[8 empty] | 1G | 0.12403 | g
[8 empty] | 1G | 0.16553 | g
[8 empty] | 1G | 0.19312 | g
[8 empty] | 1G | 0.10163 | g
[8 empty] | 1G | 0.10776 | g
[8 empty] | 1G | 0.1561 | g
[8 empty] | 1G | 4.6669 | g
[8 empty] | 1G | 5.1312 | g
[8 empty] | 1G | 5.0489 | g
[8 empty] | 1G | 5.0892 | g
[8 empty] | 1G | 5.0675 | g
[8 empty] | 1G | 5.1069 | g
[8 empty] | 1G | 5.1248 | g
[8 empty] | 1G | 5.1137 | g
[8 empty] | 1G | 5.1486 | g
[8 empty] | 1G | 5.068 | g
[8 empty] | 1G | 5.0968 | g
[8 empty] | 1G | 5.1823 | g
[8 empty] | 1G | 5.0578 | g
[8 empty] | 1G | 5.163 | g
[8 empty] | 1G | 5.1738 | g
[8 empty] | 1G | 5.1201 | g
[8 empty] | 1G | 5.1404 | g
[8 empty] | 1G | 5.0713 | g
[8 empty] | 1G | 5.1536 | g
[8 empty] | 1G | 5.1288 | g
[8 empty] | 1G | 5.1415 | g
[8 empty] | 1G | 5.151 | g
[8 empty] | 1G | 5.0897 | g
[8 empty] | 1G | 5.0795 | g
[8 empty] | 1G | 5.0939 | g
[8 empty] | 1G | 5.1658 | g
[8 empty] | 1G | 5.0972 | g
[8 empty] | 1G | 5.1477 | g
[8 empty] | 1G | 5.0725 | g
[8 empty] | 1G | 5.0701 | g
[8 empty] | 1G | 5.1665 | g
[8 empty] | 1G | 5.1295 | g
[8 empty] | 1G | 5.0958 | g
[8 empty] | 1G | 5.093 | g
[8 empty] | 1G | 5.1081 | g
[8 empty] | 1G | 5.1599 | g
[8 empty] | 1G | 5.1274 | g
[8 empty] | 1G | 5.6301 | g
[8 empty] | 1G | 0.14171 | g
[8 empty] | 1G | 0.1561 | g
[8 empty] | 1G | 0.0731 | g
[8 empty] | 1G | 0.15138 | g
[8 empty] | 1G | 0.11106 | g
[8 empty] | 1G | 0.38954 | g
[8 empty] | 1G | 9.9681 | g
[8 empty] | 1G | 10.0997 | g
[8 empty] | 1G | 10.1624 | g
[8 empty] | 1G | 10.1588 | g
[8 empty] | 1G | 10.1525 | g
[8 empty] | 1G | 10.1553 | g
[8 empty] | 1G | 10.2296 | g
[8 empty] | 1G | 10.1154 | g
[8 empty] | 1G | 10.1381 | g
[8 empty] | 1G | 10.1386 | g
[8 empty] | 1G | 10.1548 | g
[8 empty] | 1G | 10.1612 | g
[8 empty] | 1G | 10.1527 | g
[8 empty] | 1G | 10.2232 | g
[8 empty] | 1G | 10.1475 | g
[8 empty] | 1G | 10.1808 | g
[8 empty] | 1G | 10.1626 | g
[8 empty] | 1G | 10.206 | g
[8 empty] | 1G | 10.1831 | g
[8 empty] | 1G | 10.1803 | g
[8 empty] | 1G | 10.1959 | g
[8 empty] | 1G | 10.2185 | g
[8 empty] | 1G | 10.2003 | g
[8 empty] | 1G | 10.1395 | g
[8 empty] | 1G | 10.2043 | g
[8 empty] | 1G | 10.2279 | g
[8 empty] | 1G | 10.1876 | g
[8 empty] | 1G | 10.2343 | g
[8 empty] | 1G | 10.2027 | g
[8 empty] | 1G | 10.2147 | g
[8 empty] | 1G | 10.1914 | g
[8 empty] | 1G | 10.1805 | g
[8 empty] | 1G | 10.2352 | g
[8 empty] | 1G | 10.2336 | g
[8 empty] | 1G | 1.9114 | g
[8 empty] | 1G | 0.21552 | g
[8 empty] | 1G | 0.19076 | g
[8 empty] | 1G | 0.22778 | g
[8 empty] | 1G | 0.23391 | g
[8 empty] | 1G | 17.7628 | g
[8 empty] | 1G | 20.1814 | g
[8 empty] | 1G | 20.1602 | g
[8 empty] | 1G | 20.2431 | g
[8 empty] | 1G | 20.1913 | g
[8 empty] | 1G | 20.2316 | g
[8 empty] | 1G | 20.1903 | g
[8 empty] | 1G | 20.3066 | g
[8 empty] | 1G | 20.1637 | g
[8 empty] | 1G | 20.2535 | g
[8 empty] | 1G | 20.2316 | g
[8 empty] | 1G | 20.2679 | g
[8 empty] | 1G | 20.2328 | g
[8 empty] | 1G | 20.2396 | g
[8 empty] | 1G | 20.2337 | g
[8 empty] | 1G | 20.1906 | g
[8 empty] | 1G | 20.2285 | g
[8 empty] | 1G | 20.2778 | g
[8 empty] | 1G | 20.2686 | g
[8 empty] | 1G | 20.2733 | g
[8 empty] | 1G | 20.2658 | g
[8 empty] | 1G | 20.2646 | g
[8 empty] | 1G | 20.2302 | g
[8 empty] | 1G | 20.2425 | g
[8 empty] | 1G | 20.221 | g
[8 empty] | 1G | 20.2163 | g
[8 empty] | 1G | 20.2731 | g
[8 empty] | 1G | 20.2483 | g
[8 empty] | 1G | 20.2387 | g
[8 empty] | 1G | 20.1925 | g
[8 empty] | 1G | 20.3014 | g
[8 empty] | 1G | 20.295 | g
[8 empty] | 1G | 20.2321 | g
[8 empty] | 1G | 20.2201 | g
[8 empty] | 1G | 20.2891 | g
[8 empty] | 1G | 20.2865 | g
[8 empty] | 1G | 6.3172 | g
[8 empty] | 1G | 0.23132 | g
[8 empty] | 1G | 0.25395 | g
[8 empty] | 1G | 0.20137 | g
[8 empty] | 1G | 0.31809 | g
[8 empty] | 1G | 0.22495 | g
[8 empty] | 1G | 0.18133 | g
[8 empty] | 1G | 0.20349 | g
[8 empty] | 1G | 0.26904 | g
[8 empty] | 1G | 0.2358 | g
[8 empty] | 1G | 0.26221 | g
[8 empty] | 1G | 0.27117 | g
[8 empty] | 1G | 0.32304 | g
[8 empty] | 1G | 0.29875 | g
[8 empty] | 1G | 0.2806 | g
[8 empty] | 1G | 0.27777 | g
[8 empty] | 1G | 0.26645 | g
[8 empty] | 1G | 0.29333 | g
[8 empty] | 1G | 0.3129 | g
[8 empty] | 1G | 0.24122 | g
[8 empty] | 1G | 0.26338 | g
[8 empty] | 1G | 0.23061 | g
[8 empty] | 1G | 0.28414 | g
[8 empty] | 1G | 1.3155 | g
[8 empty] | 1G | 49.9894 | g
[8 empty] | 1G | 50.3011 | g
[8 empty] | 1G | 50.253 | g
[8 empty] | 1G | 50.2969 | g
[8 empty] | 1G | 50.3381 | g
[8 empty] | 1G | 50.266 | g
[8 empty] | 1G | 50.2525 | g
[8 empty] | 1G | 50.3591 | g
[8 empty] | 1G | 50.3049 | g
[8 empty] | 1G | 50.3301 | g
[8 empty] | 1G | 50.2856 | g
[8 empty] | 1G | 50.3188 | g
[8 empty] | 1G | 50.324 | g
[8 empty] | 1G | 50.2473 | g
[8 empty] | 1G | 50.3155 | g
[8 empty] | 1G | 50.299 | g
[8 empty] | 1G | 50.286 | g
[8 empty] | 1G | 50.303 | g
[8 empty] | 1G | 50.3127 | g
[8 empty] | 1G | 50.3054 | g
[8 empty] | 1G | 50.2914 | g
[8 empty] | 1G | 50.2846 | g
[8 empty] | 1G | 50.2818 | g
[8 empty] | 1G | 50.3407 | g
[8 empty] | 1G | 50.3358 | g
[8 empty] | 1G | 50.2695 | g
[8 empty] | 1G | 50.2742 | g
[8 empty] | 1G | 50.278 | g
[8 empty] | 1G | 50.3443 | g
[8 empty] | 1G | 50.3105 | g
[8 empty] | 1G | 50.2669 | g
[8 empty] | 1G | 50.2521 | g
[8 empty] | 1G | 50.2146 | g
[8 empty] | 1G | 50.2521 | g
[8 empty] | 1G | 50.3589 | g
[8 empty] | 1G | 50.2815 | g
[8 empty] | 1G | 50.3013 | g
[8 empty] | 1G | 50.2903 | g
[8 empty] | 1G | 50.324 | g
[8 empty] | 1G | 50.3282 | g
[8 empty] | 1G | 50.9118 | g
[8 empty] | 1G | 15.7005 | g
[8 empty] | 1G | 0.32068 | g
[8 empty] | 1G | 0.30253 | g
[8 empty] | 1G | 0.26173 | g
[8 empty] | 1G | 0.27187 | g
[8 empty] | 1G | 0.27918 | g
[8 empty] | 1G | 0.29262 | g
[8 empty] | 1G | 0.22943 | g
[8 empty] | 1G | 0.30088 | g
[8 empty] | 1G | 0.29404 | g
[8 empty] | 1G | 0.30489 | g
[8 empty] | 1G | 0.26386 | g
[8 empty] | 1G | 0.30276 | g
[8 empty] | 1G | 42.9318 | g
[8 empty] | 1G | 100.2841 | g
[8 empty] | 1G | 100.2903 | g
[8 empty] | 1G | 100.2511 | g
[8 empty] | 1G | 100.2575 | g
[8 empty] | 1G | 100.2568 | g
[8 empty] | 1G | 100.2504 | g
[8 empty] | 1G | 100.3292 | g
[8 empty] | 1G | 100.2839 | g
[8 empty] | 1G | 100.2827 | g
[8 empty] | 1G | 100.2438 | g
[8 empty] | 1G | 100.2665 | g
[8 empty] | 1G | 100.3197 | g
[8 empty] | 1G | 100.3495 | g
[8 empty] | 1G | 100.2683 | g
[8 empty] | 1G | 100.3426 | g
[8 empty] | 1G | 100.3047 | g
[8 empty] | 1G | 100.3535 | g
[8 empty] | 1G | 100.287 | g
[8 empty] | 1G | 100.3436 | g
[8 empty] | 1G | 100.3054 | g
[8 empty] | 1G | 100.2785 | g
[8 empty] | 1G | 100.3174 | g
[8 empty] | 1G | 100.2844 | g
[8 empty] | 1G | 100.2233 | g
[8 empty] | 1G | 100.2912 | g
[8 empty] | 1G | 100.3629 | g
[8 empty] | 1G | 100.3089 | g
[8 empty] | 1G | 100.3223 | g
[8 empty] | 1G | 100.3313 | g
[8 empty] | 1G | 100.3339 | g
[8 empty] | 1G | 100.3554 | g
[8 empty] | 1G | 100.2924 | g
[8 empty] | 1G | 100.2801 | g
[8 empty] | 1G | 100.2985 | g
[8 empty] | 1G | 100.3322 | g
[8 empty] | 1G | 100.3367 | g
[8 empty] | 1G | 100.2792 | g
[8 empty] | 1G | 100.2646 | g
[8 empty] | 1G | 100.2417 | g
[8 empty] | 1G | 101.2677 | g
[8 empty] | 1G | 39.9244 | g
[8 empty] | 1G | 0.33743 | g
[8 empty] | 1G | 0.31149 | g
[8 empty] | 1G | 0.3511 | g
[8 empty] | 1G | 0.30394 | g
^Jz
[8 empty] | 1G | 0.39567 | g
^J
[8 empty] | Z1G | -0.08017 | g
r
[8 empty] | 1G | 0.0099 | g
[8 empty] | 1G | 0.03325 | g
[8 empty] | 1G | 0.03773 | g
[8 empty] | 1G | 0.10941 | g
[8 empty] | 1G | 0.04622 | g
[8 empty] | 1G | 0.06107 | g
[8 empty] | 1G | 0.09243 | g
[8 empty] | 1G | 0.10446 | g
[8 empty] | 1G | 0.07333 | g
[8 empty] | 1G | 0.06767 | g
[8 empty] | 1G | 0.05282 | g
[8 empty] | 1G | 0.06838 | g
[8 empty] | 1G | 0.93564 | g
[8 empty] | 1G | 1.0356 | g
[8 empty] | 1G | 1.1655 | g
[8 empty] | 1G | 1.1047 | g
[8 empty] | 1G | 1.1049 | g
[8 empty] | 1G | 1.0554 | g
[8 empty] | 1G | 1.0498 | g
[8 empty] | 1G | 1.1464 | g
[8 empty] | 1G | 1.1094 | g
[8 empty] | 1G | 1.157 | g
[8 empty] | 1G | 1.0554 | g
[8 empty] | 1G | 1.1057 | g
[8 empty] | 1G | 1.1212 | g
[8 empty] | 1G | 1.1224 | g
[8 empty] | 1G | 1.0953 | g
[8 empty] | 1G | 1.1698 | g
[8 empty] | 1G | 1.1196 | g
[8 empty] | 1G | 1.1014 | g
[8 empty] | 1G | 1.1198 | g
[8 empty] | 1G | 1.1809 | g
[8 empty] | 1G | 1.1009 | g
[8 empty] | 1G | 1.1259 | g
[8 empty] | 1G | 1.145 | g
[8 empty] | 1G | 1.1233 | g
[8 empty] | 1G | 1.109 | g
[8 empty] | 1G | 1.1745 | g
[8 empty] | 1G | 1.1224 | g
[8 empty] | 1G | 1.1174 | g
[8 empty] | 1G | 1.1219 | g
[8 empty] | 1G | 1.1097 | g
[8 empty] | 1G | 1.1 | g
[8 empty] | 1G | 1.1033 | g
[8 empty] | 1G | 1.1587 | g
[8 empty] | 1G | 1.0887 | g
[8 empty] | 1G | 1.0995 | g
[8 empty] | 1G | 1.0865 | g
[8 empty] | 1G | 1.1177 | g
[8 empty] | 1G | 1.1203 | g
[8 empty] | 1G | 0.21693 | g
[8 empty] | 1G | 0.06885 | g
[8 empty] | 1G | 0.13818 | g
[8 empty] | 1G | 0.13983 | g
[8 empty] | 1G | 0.13346 | g
[8 empty] | 1G | 0.09974 | g
[8 empty] | 1G | 0.11082 | g
[8 empty] | 1G | 0.08536 | g
[8 empty] | 1G | 0.0922 | g
[8 empty] | 1G | 0.12497 | g
[8 empty] | 1G | 0.06272 | g
[8 empty] | 1G | 0.12662 | g
[8 empty] | 1G | 1.0302 | g
[8 empty] | 1G | 2.1486 | g
[8 empty] | 1G | 2.1259 | g
[8 empty] | 1G | 2.1479 | g
[8 empty] | 1G | 2.1118 | g
[8 empty] | 1G | 2.0606 | g
[8 empty] | 1G | 2.1224 | g
[8 empty] | 1G | 2.2012 | g
[8 empty] | 1G | 2.1571 | g
[8 empty] | 1G | 2.1156 | g
[8 empty] | 1G | 2.1175 | g
[8 empty] | 1G | 2.1422 | g
[8 empty] | 1G | 2.1674 | g
[8 empty] | 1G | 2.1377 | g
[8 empty] | 1G | 2.1946 | g
[8 empty] | 1G | 2.1644 | g
[8 empty] | 1G | 2.1538 | g
[8 empty] | 1G | 2.105 | g
[8 empty] | 1G | 2.1561 | g
[8 empty] | 1G | 2.1205 | g
[8 empty] | 1G | 2.1373 | g
[8 empty] | 1G | 2.1446 | g
[8 empty] | 1G | 2.1156 | g
[8 empty] | 1G | 2.1278 | g
[8 empty] | 1G | 2.1547 | g
[8 empty] | 1G | 2.1561 | g
[8 empty] | 1G | 2.1149 | g
[8 empty] | 1G | 2.1292 | g
[8 empty] | 1G | 2.1486 | g
[8 empty] | 1G | 2.1818 | g
[8 empty] | 1G | 2.1545 | g
[8 empty] | 1G | 2.1141 | g
[8 empty] | 1G | 2.1729 | g
[8 empty] | 1G | 2.1226 | g
[8 empty] | 1G | 4.1642 | g
[8 empty] | 1G | 1.4018 | g
[8 empty] | 1G | 0.1436 | g
[8 empty] | 1G | 0.11012 | g
[8 empty] | 1G | 0.1311 | g
^Jz
[8 empty] | 1G | 0.0896 | g
[8 empty] | 1G | 0.11648 | g
^J
[8 empty] | Z1G | -0.13912 | g
r
[8 empty] | 1G | -0.08371 | g
[8 empty] | 1G | -0.09455 | g
[8 empty] | 1G | -0.09361 | g
[8 empty] | 1G | -0.07404 | g
[8 empty] | 1G | -0.11106 | g
[8 empty] | 1G | -0.08607 | g
[8 empty] | 1G | -0.09455 | g
[8 empty] | 1G | -0.01108 | g
[8 empty] | 1G | -0.04032 | g
[8 empty] | 1G | -0.05353 | g
[8 empty] | 1G | 3.2559 | g
[8 empty] | 1G | 4.9527 | g
[8 empty] | 1G | 4.9258 | g
[8 empty] | 1G | 4.9661 | g
[8 empty] | 1G | 4.9053 | g
[8 empty] | 1G | 4.9145 | g
[8 empty] | 1G | 4.9317 | g
[8 empty] | 1G | 4.8944 | g
[8 empty] | 1G | 4.9597 | g
[8 empty] | 1G | 4.9546 | g
[8 empty] | 1G | 4.8819 | g
[8 empty] | 1G | 4.9157 | g
[8 empty] | 1G | 4.9781 | g
[8 empty] | 1G | 4.9374 | g
[8 empty] | 1G | 4.9175 | g
[8 empty] | 1G | 4.981 | g
[8 empty] | 1G | 4.9961 | g
[8 empty] | 1G | 4.9232 | g
[8 empty] | 1G | 4.9767 | g
[8 empty] | 1G | 4.9786 | g
[8 empty] | 1G | 4.9635 | g
[8 empty] | 1G | 4.9737 | g
[8 empty] | 1G | 4.9852 | g
[8 empty] | 1G | 4.9649 | g
[8 empty] | 1G | 4.9817 | g
[8 empty] | 1G | 4.9715 | g
[8 empty] | 1G | 4.9885 | g
[8 empty] | 1G | 4.9916 | g
[8 empty] | 1G | 4.8996 | g
[8 empty] | 1G | 4.9902 | g
[8 empty] | 1G | 4.9425 | g
[8 empty] | 1G | 5.0121 | g
[8 empty] | 1G | 4.9668 | g
[8 empty] | 1G | 4.9317 | g
[8 empty] | 1G | 4.9607 | g
[8 empty] | 1G | 4.9673 | g
[8 empty] | 1G | 4.9635 | g
[8 empty] | 1G | 4.9204 | g
[8 empty] | 1G | 4.9692 | g
[8 empty] | 1G | 4.9249 | g
[8 empty] | 1G | 0.74535 | g
[8 empty] | 1G | 0.01981 | g
[8 empty] | 1G | -0.06296 | g
[8 empty] | 1G | -0.05022 | g
[8 empty] | 1G | -0.00613 | g
[8 empty] | 1G | -0.0764 | g
[8 empty] | 1G | -0.00141 | g
[8 empty] | 1G | -0.05588 | g
[8 empty] | 1G | -0.02523 | g
[8 empty] | 1G | 9.3956 | g
[8 empty] | 1G | 9.9568 | g
[8 empty] | 1G | 10.0167 | g
[8 empty] | 1G | 10.0159 | g
[8 empty] | 1G | 9.9061 | g
[8 empty] | 1G | 9.9992 | g
[8 empty] | 1G | 9.9391 | g
[8 empty] | 1G | 9.9933 | g
[8 empty] | 1G | 10.0077 | g
[8 empty] | 1G | 10.0181 | g
[8 empty] | 1G | 9.9259 | g
[8 empty] | 1G | 10.0195 | g
[8 empty] | 1G | 9.9313 | g
[8 empty] | 1G | 9.9662 | g
[8 empty] | 1G | 10.0051 | g
[8 empty] | 1G | 10.0256 | g
[8 empty] | 1G | 9.9556 | g
[8 empty] | 1G | 9.9424 | g
[8 empty] | 1G | 9.9916 | g
[8 empty] | 1G | 9.9485 | g
[8 empty] | 1G | 10.0103 | g
[8 empty] | 1G | 9.9674 | g
[8 empty] | 1G | 10.0117 | g
[8 empty] | 1G | 9.9676 | g
[8 empty] | 1G | 10.0256 | g
[8 empty] | 1G | 9.9457 | g
[8 empty] | 1G | 9.9881 | g
[8 empty] | 1G | 10.0508 | g
[8 empty] | 1G | 9.9282 | g
[8 empty] | 1G | 9.9535 | g
[8 empty] | 1G | 9.9742 | g
[8 empty] | 1G | 10.0041 | g
[8 empty] | 1G | 10.014 | g
[8 empty] | 1G | 9.9742 | g
[8 empty] | 1G | 9.9164 | g
[8 empty] | 1G | 9.9374 | g
[8 empty] | 1G | 9.4229 | g
[8 empty] | 1G | 0.00165 | g
[8 empty] | 1G | -0.04268 | g
[8 empty] | 1G | -0.00472 | g
[8 empty] | 1G | 0.02712 | g
[8 empty] | 1G | 0.0316 | g
[8 empty] | 1G | -0.00094 | g
[8 empty] | 1G | 9.4349 | g
[8 empty] | 1G | 19.9767 | g
[8 empty] | 1G | 20.0505 | g
[8 empty] | 1G | 20.0276 | g
[8 empty] | 1G | 19.9649 | g
[8 empty] | 1G | 19.9555 | g
[8 empty] | 1G | 19.9725 | g
[8 empty] | 1G | 19.9963 | g
[8 empty] | 1G | 19.9826 | g
[8 empty] | 1G | 20.039 | g
[8 empty] | 1G | 19.9793 | g
[8 empty] | 1G | 20.0001 | g
[8 empty] | 1G | 20.0107 | g
[8 empty] | 1G | 20.0257 | g
[8 empty] | 1G | 20.0481 | g
[8 empty] | 1G | 20.0104 | g
[8 empty] | 1G | 20.0364 | g
[8 empty] | 1G | 20.0529 | g
[8 empty] | 1G | 20.0135 | g
[8 empty] | 1G | 19.9621 | g
[8 empty] | 1G | 20.0015 | g
[8 empty] | 1G | 20.0229 | g
[8 empty] | 1G | 20.0406 | g
[8 empty] | 1G | 20.0222 | g
[8 empty] | 1G | 20.0533 | g
[8 empty] | 1G | 20.0177 | g
[8 empty] | 1G | 20.0291 | g
[8 empty] | 1G | 20.0543 | g
[8 empty] | 1G | 20.1003 | g
[8 empty] | 1G | 20.0305 | g
[8 empty] | 1G | 20.1052 | g
[8 empty] | 1G | 20.0904 | g
[8 empty] | 1G | 20.0041 | g
[8 empty] | 1G | 20.0698 | g
[8 empty] | 1G | 19.9939 | g
[8 empty] | 1G | 20.0576 | g
[8 empty] | 1G | 19.9934 | g
[8 empty] | 1G | 20.0274 | g
[8 empty] | 1G | 20.034 | g
[8 empty] | 1G | 20.0489 | g
[8 empty] | 1G | 20.0437 | g
[8 empty] | 1G | 20.0491 | g
[8 empty] | 1G | 20.0321 | g
[8 empty] | 1G | 19.9989 | g
[8 empty] | 1G | 20.0951 | g
[8 empty] | 1G | 20.0581 | g
[8 empty] | 1G | 20.0128 | g
[8 empty] | 1G | 20.0288 | g
[8 empty] | 1G | 13.6562 | g
[8 empty] | 1G | 0.0448 | g
[8 empty] | 1G | 0.01863 | g
[8 empty] | 1G | 0.01226 | g
[8 empty] | 1G | 0.04362 | g
[8 empty] | 1G | 0.1469 | g
[8 empty] | 1G | 0.07758 | g
[8 empty] | 1G | 0.06178 | g
[8 empty] | 1G | 0.11177 | g
[8 empty] | 1G | 0.05022 | g
[8 empty] | 1G | 0.15515 | g
[8 empty] | 1G | 30.0235 | g
[8 empty] | 1G | 50.1391 | g
[8 empty] | 1G | 50.0656 | g
[8 empty] | 1G | 50.1108 | g
[8 empty] | 1G | 50.0979 | g
[8 empty] | 1G | 50.0309 | g
[8 empty] | 1G | 50.0339 | g
[8 empty] | 1G | 50.1071 | g
[8 empty] | 1G | 50.0707 | g
[8 empty] | 1G | 50.0137 | g
[8 empty] | 1G | 50.0585 | g
[8 empty] | 1G | 50.0276 | g
[8 empty] | 1G | 50.0903 | g
[8 empty] | 1G | 50.0981 | g
[8 empty] | 1G | 50.0771 | g
[8 empty] | 1G | 50.1101 | g
[8 empty] | 1G | 50.0983 | g
[8 empty] | 1G | 50.1054 | g
[8 empty] | 1G | 50.1005 | g
[8 empty] | 1G | 50.0064 | g
[8 empty] | 1G | 50.0382 | g
[8 empty] | 1G | 50.1193 | g
[8 empty] | 1G | 50.0738 | g
[8 empty] | 1G | 50.0212 | g
[8 empty] | 1G | 50.0641 | g
[8 empty] | 1G | 50.1346 | g
[8 empty] | 1G | 50.1158 | g
[8 empty] | 1G | 50.062 | g
[8 empty] | 1G | 50.0726 | g
[8 empty] | 1G | 50.0806 | g
[8 empty] | 1G | 50.1049 | g
[8 empty] | 1G | 50.1071 | g
[8 empty] | 1G | 50.1012 | g
[8 empty] | 1G | 50.0698 | g
[8 empty] | 1G | 50.1037 | g
[8 empty] | 1G | 50.0604 | g
[8 empty] | 1G | 50.045 | g
[8 empty] | 1G | 50.016 | g
[8 empty] | 1G | 50.0592 | g
[8 empty] | 1G | 50.1243 | g
[8 empty] | 1G | 50.0986 | g
[8 empty] | 1G | 50.1184 | g
[8 empty] | 1G | 50.1471 | g
[8 empty] | 1G | 50.0597 | g
[8 empty] | 1G | 50.1002 | g
[8 empty] | 1G | 50.0424 | g
[8 empty] | 1G | 33.4334 | g
[8 empty] | 1G | 0.08748 | g
[8 empty] | 1G | 0.0547 | g
[8 empty] | 1G | 0.05918 | g
[8 empty] | 1G | 0.03867 | g
[8 empty] | 1G | 0.00825 | g
[8 empty] | 1G | 0.06626 | g
[8 empty] | 1G | 0.07498 | g
[8 empty] | 1G | 0.08371 | g
[8 empty] | 1G | 0.10422 | g
[8 empty] | 1G | 1.2563 | g
[8 empty] | 1G | 97.8429 | g
[8 empty] | 1G | 99.9613 | g
[8 empty] | 1G | 100.0877 | g
[8 empty] | 1G | 100.083 | g
[8 empty] | 1G | 100.0483 | g
[8 empty] | 1G | 99.9962 | g
[8 empty] | 1G | 100.046 | g
[8 empty] | 1G | 100.0832 | g
[8 empty] | 1G | 100.0797 | g
[8 empty] | 1G | 100.0828 | g
[8 empty] | 1G | 100.0738 | g
[8 empty] | 1G | 100.038 | g
[8 empty] | 1G | 100.0813 | g
[8 empty] | 1G | 100.0049 | g
[8 empty] | 1G | 100.0509 | g
[8 empty] | 1G | 100.0186 | g
[8 empty] | 1G | 100.0233 | g
[8 empty] | 1G | 100.0476 | g
[8 empty] | 1G | 100.0568 | g
[8 empty] | 1G | 100.0483 | g
[8 empty] | 1G | 100.074 | g
[8 empty] | 1G | 100.0519 | g
[8 empty] | 1G | 100.0724 | g
[8 empty] | 1G | 100.0788 | g
[8 empty] | 1G | 100.0406 | g
[8 empty] | 1G | 100.1382 | g
[8 empty] | 1G | 100.0946 | g
[8 empty] | 1G | 100.0776 | g
[8 empty] | 1G | 100.0502 | g
[8 empty] | 1G | 100.0632 | g
[8 empty] | 1G | 100.1092 | g
[8 empty] | 1G | 100.0993 | g
[8 empty] | 1G | 100.0521 | g
[8 empty] | 1G | 100.146 | g
[8 empty] | 1G | 100.1082 | g
[8 empty] | 1G | 100.0792 | g
[8 empty] | 1G | 100.0842 | g
[8 empty] | 1G | 100.0986 | g
[8 empty] | 1G | 100.1469 | g
[8 empty] | 1G | 100.0658 | g
[8 empty] | 1G | 100.1181 | g
[8 empty] | 1G | 100.0689 | g
[8 empty] | 1G | 100.038 | g
[8 empty] | 1G | 100.0809 | g
[8 empty] | 1G | 100.1012 | g
[8 empty] | 1G | 47.4355 | g
[8 empty] | 1G | 0.14737 | g
[8 empty] | 1G | 0.1179 | g
[8 empty] | 1G | 0.17025 | g
[8 empty] | 1G | 0.09125 | g
[8 empty] | 1G | 0.12969 | g
[8 empty] | 1G | 0.11554 | g
[8 empty] | 1G | 0.12733 | g
[8 empty] | 1G | 0.14478 | g
[8 empty] | 1G | 0.13605 | g
[8 empty] | 1G | 0.12332 | g
[8 empty] | 1G | 0.0731 | g
[8 empty] | 1G | 0.12497 | g
[8 empty] | 1G | 0.09668 | g
[8 empty] | 1G | 0.08984 | g
[8 empty] | 1G | 0.12827 | g
[8 empty] | 1G | 0.11696 | g
[8 empty] | 1G | 0.17284 | g

## 200414 - 100g vs WAOAW and 500g
Mass Balance Comparison Test
| Device Under Test: |  | OS_Nano_Balance (3D printed balance), Using TAL221 100g load cell
| Lab Balance: |  | Comparing with TAL221 500g loadcell and WAOAW 500g/0.01g scale
| DUT Calibration: |  | WAOAW measured tupperware with oatmeal to be 74.96g
|  |  | 3207.887 | g/div
Empty container |  |  |  |  | Container w/ sand |  |  |  | Container w/sand | [15 empty] | Scale upper limit
| WAOAW | 500g | 100g |  | WAOAW | 500g | 100g |  | WAOAW | 500g | 100g |  | WAOAW | 500g | 100g | WAOAW | 500g | 100g | APX402 | DUT | APX402 | DUT | APX402 | DUT | APX402 | DUT
| 24.99 | 24.98 | 24.971 |  | 50.03 | 50.02 | 49.998 |  | 75 | 74.97 | 74.957 |  | 100.03 | 100.02 | 100.019
| 24.98 | 24.92 | 24.967 |  | 50.04 | 50.02 | 49.999 |  | 74.98 | 74.92 | 74.959 |  | 100.03 | 100 | 100.019
| 24.99 | 24.94 | 24.966 |  | 50.04 | 49.98 | 50.001 |  | 74.99 | 74.87 | 74.95 |  | 100.03 | 100.01 | 100.027
| 24.97 | 24.96 | 24.96 |  | 50.04 | 50 | 50.006 |  | 74.97 | 74.96 | 74.965 |  | 100.03 | 100.01 | 100.035
| 24.99 | 24.9 | 24.972 |  | 50.04 | 49.96 | 50.001 |  | 74.99 | 74.94 | 74.959 |  | 100.04 | 99.98 | 100.026
Average | 24.984 | 24.94 |  |  | 50.038 | 49.996 |  |  | 74.986 | 74.932 |  |  | 100.032 | 100.004 |  | #DIV/0! | #DIV/0!
| Notes |  |  |  | Notes |  |  |  | Notes |  |  |  | Notes |  |  | Notes
Wiring is the issue - need to flip it around; mine is upside-down and the wires don't reach; will resolve these issues and re-test next week.
| WAOAW | 500g | 100g |  | Average
| 24.99 | 24.98 | 24.971 |  | 24.984 | 50.038 | 74.986 | 100.032 | 0 | 0 | 0 | 0
| 24.98 | 24.92 | 24.967 |  | 24.94 | 49.996 | 74.932 | 100.004 | 0 | nan | nan | nan
| 24.99 | 24.94 | 24.966 | Tare | 24.9672 | 50.001 | 74.958 | 100.0252 | 0 | 0 | 0 | 0
| 24.97 | 24.96 | 24.96 |  | Stdev
| 24.99 | 24.9 | 24.972 |  | 0.0089443 | 0.0044721 | 0.011402 | 0.0044721 | 0 | 0 | 0 | 0
| 50.03 | 50.02 | 49.998 |  | 0.031623 | 0.026077 | 0.039623 | 0.015166 | 0 | nan | nan | nan
| 50.04 | 50.02 | 49.999 |  | 0.0047645 | 0.0030822 | 0.0053852 | 0.0066483 | 0 | 0 | 0 | 0
| 50.04 | 49.98 | 50.001 |  | max
| 50.04 | 50 | 50.006 |  | 24.99 | 50.04 | 75 | 100.04 | 0 | 0 | 0 | 0
| 50.04 | 49.96 | 50.001 |  | 24.98 | 50.02 | 74.97 | 100.02 | 0 | nan | nan | nan
| 75 | 74.97 | 74.957 |  | 24.972 | 50.006 | 74.965 | 100.035 | 0 | 0 | 0 | 0
| 74.98 | 74.92 | 74.959 |  | min
| 74.99 | 74.87 | 74.95 |  | 24.97 | 50.03 | 74.97 | 100.03 | 0 | 0 | 0 | 0
| 74.97 | 74.96 | 74.965 |  | 24.9 | 49.96 | 74.87 | 99.98 | 0 | nan | nan | nan
| 74.99 | 74.94 | 74.959 |  | 24.96 | 49.998 | 74.95 | 100.019 | 0 | 0 | 0 | 0
| 100.03 | 100.02 | 100.019
| 100.03 | 100 | 100.019
| 100.03 | 100.01 | 100.027
| 100.03 | 100.01 | 100.035
| 100.04 | 99.98 | 100.026
| 0 | 0 | 0
| 0 | 0 | 0
| 0 | 0 | 0
| 0 | 0 | 0
| 0 | 0 | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0 | Tare
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0
| 0 |  | 0

## 200323 - 500g vs A-160
Mass Balance Comparison Test
| Device Under Test: |  | OS_Nano_Balance (3D printed balance), Using TAL221 500g load cell
| Lab Balance: |  | Denver A-160 160/0.1mg scale, S/N 78645 Calibrated 12/19/2019 by B Kniffin
| DUT Calibration: |  | A-160 measured plastic container to be 16.4338g
|  |  | 3584.2406 | g/div
Empty container |  |  |  |  | Container w/ sand |  |  |  | Container w/sand | [16 empty] | Scale upper limit
| APX402 (g) | A-160 (g) | DUT (g) |  | APX402 (g) | A-160 (g) | DUT (g) |  | APX402 (g) | A-160 (g) | DUT (g) |  | APX402 (g) | A-160 (g) | DUT (g) | APX402 (g) | A-160 (g) | DUT (g) | APX402 | DUT | APX402 | DUT |  | APX402 | DUT | APX402 | DUT
| 16.4 | 16.4311 | 16.438 |  | 25.02 | 25.0462 | 25.033 |  | 50.05 | 50.0808 | 50.068 |  | 100.03 | 100.0353 | 100.055 | 150 | 149.9747 | 149.9678 | 200.07 | 200.104 | 250.07 | 250.145 |  | 301.14 | 301.166 | 332.89 | 332.917
| 16.4 | 16.4292 | 16.406 |  | 25.02 | 25.0455 | 25.032 |  | 50.06 | 50.0746 | 50.066 |  | 100.03 | 100.0342 | 100.033 | 150 | 149.9721 | 149.983 | 200.08 | 200.121 | 250.07 | 250.029 | Tare | 301.13 | 301.106 | 332.89 | 332.931
| 16.4 | 16.4323 | 16.424 | Tare | 25.02 | 25.0448 | 25.036 |  | 50.05 | 50.0747 | 50.075 |  | 100.04 | 100.0339 | 100.048 | 149.99 | 149.9723 | 150.006 | 200.09 | 200.103 | 250.06 | 250.046 |  | 301.13 | 301.163 | 332.88 | 332.916
| 16.4 | 16.4313 | 16.422 |  | 25.02 | 25.0418 | 25.044 |  | 50.05 | 50.0773 | 50.086 |  | 100.03 | 100.0484 | 100.067 | 150.01 | 149.9723 | 150.012 | 200.08 | 200.114 | 250.06 | 250.042 |  | 301.13 | 301.136 | 332.88 | 332.949
| 16.4 | 16.4331 | 16.403 |  | 25.02 | 25.0423 | 25.029 |  | 50.05 | 50.0732 | 50.12 |  | 100.03 | 100.036 | 100.068 | 150 | 149.9719 | 150.022 | 200.07 | 200.125 | 250.07 | 250.012 |  | 301.13 | 301.149 | 332.89 | 332.926
Average | 16.4 | 16.4314 |  |  | 25.02 | 25.0441 |  |  | 50.052 | 50.0761 |  |  | 100.032 | 100.0376 |  | 150 | 149.9727
| Notes |  |  |  | Notes |  |  |  | Notes |  |  |  | Notes |  |  | Notes
Wiring is the issue - need to flip it around; mine is upside-down and the wires don't reach; will resolve these issues and re-test next week.
| APX402 (g) | A-160 (g) | DUT (g) |  | Average
| 16.4 | 16.4311 | 16.438 |  | 16.4 | 25.02 | 50.052 | 100.032 | 150 | 200.078 | 250.066 | 301.132
| 16.4 | 16.4292 | 16.406 |  | 16.4314 | 25.0441 | 50.0761 | 100.0376 | 149.9727 | nan | nan | nan
| 16.4 | 16.4323 | 16.424 | Tare | 16.4186 | 25.0348 | 50.083 | 100.0542 | 149.9982 | 200.1134 | 250.0548 | 301.144
| 16.4 | 16.4313 | 16.422 |  | Stdev
| 16.4 | 16.4331 | 16.403 |  | 0 | 0 | 0.0044721 | 0.0044721 | 0.0070711 | 0.0083666 | 0.0054772 | 0.0044721
| 25.02 | 25.0462 | 25.033 |  | 0.0014697 | 0.0019614 | 0.0030062 | 0.0061183 | 0.0011524 | nan | nan | nan
| 25.02 | 25.0455 | 25.032 |  | 0.014311 | 0.0057184 | 0.022113 | 0.014516 | 0.022209 | 0.0098641 | 0.052141 | 0.024382
| 25.02 | 25.0448 | 25.036 |  | max
| 25.02 | 25.0418 | 25.044 |  | 16.4 | 25.02 | 50.06 | 100.04 | 150.01 | 200.09 | 250.07 | 301.14
| 25.02 | 25.0423 | 25.029 |  | 16.4331 | 25.0462 | 50.0808 | 100.0484 | 149.9747 | nan | nan | nan
| 50.05 | 50.0808 | 50.068 |  | 16.438 | 25.044 | 50.12 | 100.068 | 150.022 | 200.125 | 250.145 | 301.166
| 50.06 | 50.0746 | 50.066 |  | min
| 50.05 | 50.0747 | 50.075 |  | 16.4 | 25.02 | 50.05 | 100.03 | 149.99 | 200.07 | 250.06 | 301.13
| 50.05 | 50.0773 | 50.086 |  | 16.4292 | 25.0418 | 50.0732 | 100.0339 | 149.9719 | nan | nan | nan
| 50.05 | 50.0732 | 50.12 |  | 16.403 | 25.029 | 50.066 | 100.033 | 149.9678 | 200.103 | 250.012 | 301.106
| 100.03 | 100.0353 | 100.055
| 100.03 | 100.0342 | 100.033
| 100.04 | 100.0339 | 100.048
| 100.03 | 100.0484 | 100.067
| 100.03 | 100.036 | 100.068
| 150 | 149.9747 | 149.9678
| 150 | 149.9721 | 149.983
| 149.99 | 149.9723 | 150.006
| 150.01 | 149.9723 | 150.012
| 150 | 149.9719 | 150.022
| 200.07 |  | 200.104
| 200.08 |  | 200.121
| 200.09 |  | 200.103
| 200.08 |  | 200.114
| 200.07 |  | 200.125
| 250.07 |  | 250.145
| 250.07 |  | 250.029 | Tare
| 250.06 |  | 250.046
| 250.06 |  | 250.042
| 250.07 |  | 250.012
| 301.14 |  | 301.166
| 301.13 |  | 301.106
| 301.13 |  | 301.163
| 301.13 |  | 301.136
| 301.13 |  | 301.149
| 332.89 |  | 332.917
| 332.89 |  | 332.931
| 332.88 |  | 332.916
| 332.88 |  | 332.949
| 332.89 |  | 332.926

## 200221 - 5000g vs A-160 LC1
Mass Balance Comparison Test
| Device Under Test: |  | OS_Nano_Balance (3D printed balance), Using TAL221 500g load cell
| Lab Balance: |  | Denver A-160 160/0.1mg scale, S/N 78645 Calibrated 12/19/2019 by B Kniffin
| DUT Calibration: |  | A-160 measured plastic container with sand to be 99.9907g
423.70242 div/g
Empty container |  |  |  |  | Container w/ sand |  |  |  | Container w/sand
| APX402 (g) | A-160 (g) | DUT (g) |  | APX402 (g) | A-160 (g) | DUT (g) |  | APX402 (g) | A-160 (g) | DUT (g) |  | APX402 (g) | A-160 (g) | DUT (g) | APX402 (g) | A-160 (g) | DUT (g) | APX402 | DUT
| 16.38 | 16.3827 | 16.35 |  | 30 | 30.0031 | 29.98 |  | 60 | 60.0025 | 59.95 |  | 100 | 99.9854 | 99.96 | 150 | 149.9755 | 149.97 | 200 | 199.9 |  | 300 | 299.89
| 16.38 | 16.3822 | 16.36 |  | 30 | 30.0031 | 29.95 |  | 60 | 60.0025 | 59.97 |  | 99.99 | 99.9851 | 99.95 | 150 | 149.9751 | 149.92 | 200 | 199.97 | tare | 299.99 | 299.86
| 16.38 | 16.3821 | 16.36 |  | 30 | 30.0031 | 29.94 |  | 60 | 60.0026 | 59.95 |  | 99.99 | 99.9851 | 99.95 | 150 | 149.9751 | 149.96 | 200 | 199.96 |  | 299.99 | 299.88
| 16.38 | 16.3819 | 16.34 |  | 30 | 30.0028 | 29.95 |  | 60 | 60.0025 | 59.97 |  | 99.99 | 99.9849 | 99.94 | 150.01 | 149.9751 | 149.93 | 200.01 | 199.96 |  | 299.99 | 299.88
| 16.37 | 16.3818 | 16.36 |  | 30 | 30.0026 | 29.98 |  | 60 | 60.0021 | 59.98 |  | 100 | 99.985 | 99.93 | 150.01 | 149.9747 | 149.93 | 200 | 199.93 |  | 299.99 | 299.88
Average | 16.378 | 16.3821 |  |  | 30 | 30.0029 |  |  | 60 | 60.0024 |  |  | 99.994 | 99.9851 |  | 150.004 | 149.9751
| Notes |  |  |  | Notes |  |  |  | Notes |  |  |  | Notes |  |  | Notes
Wiring is the issue - need to flip it around; mine is upside-down and the wires don't reach; will resolve these issues and re-test next week.
| APX402 (g) | A-160 (g) | DUT (g) |  | Average
| 16.38 | 16.3827 | 16.35 |  | 16.378 | 30 | 60 | 99.994 | 150.004 | 200.002 | 299.992 | nan
| 16.38 | 16.3822 | 16.36 |  | 16.3821 | 30.0029 | 60.0024 | 99.9851 | 149.9751 | nan | nan | nan
| 16.38 | 16.3821 | 16.36 | Tare | 16.354 | 29.96 | 59.964 | 99.946 | 149.942 | 199.944 | 299.878 | nan
| 16.38 | 16.3819 | 16.34 |  | Stdev
| 16.37 | 16.3818 | 16.36 |  | 0.0044721 | 0 | 0 | 0.0054772 | 0.0054772 | 0.0044721 | 0.0044721 | nan
| 30 | 30.0031 | 29.98 |  | 0.00035071 | 0.00023022 | 0.00019494 | 0.00018708 | 0.00028284 | nan | nan | nan
| 30 | 30.0031 | 29.95 |  | 0.0089443 | 0.018708 | 0.013416 | 0.011402 | 0.021679 | 0.02881 | 0.010954 | nan
| 30 | 30.0031 | 29.94 |  | max
| 30 | 30.0028 | 29.95 |  | 16.38 | 30 | 60 | 100 | 150.01 | 200.01 | 300 | nan
| 30 | 30.0026 | 29.98 |  | 16.3827 | 30.0031 | 60.0026 | 99.9854 | 149.9755 | nan | nan | nan
| 60 | 60.0025 | 59.95 |  | 16.36 | 29.98 | 59.98 | 99.96 | 149.97 | 199.97 | 299.89 | nan
| 60 | 60.0025 | 59.97 |  | min
| 60 | 60.0026 | 59.95 |  | 16.37 | 30 | 60 | 99.99 | 150 | 200 | 299.99 | nan
| 60 | 60.0025 | 59.97 |  | 16.3818 | 30.0026 | 60.0021 | 99.9849 | 149.9747 | nan | nan | nan
| 60 | 60.0021 | 59.98 |  | 16.34 | 29.94 | 59.95 | 99.93 | 149.92 | 199.9 | 299.86 | nan
| 100 | 99.9854 | 99.96
| 99.99 | 99.9851 | 99.95
| 99.99 | 99.9851 | 99.95
| 99.99 | 99.9849 | 99.94
| 100 | 99.985 | 99.93
| 150 | 149.9755 | 149.97
| 150 | 149.9751 | 149.92
| 150 | 149.9751 | 149.96
| 150.01 | 149.9751 | 149.93
| 150.01 | 149.9747 | 149.93
| 200 |  | 199.9
| 200 |  | 199.97
| 200 |  | 199.96
| 200.01 |  | 199.96
| 200 |  | 199.93
| 300 |  | 299.89
| 299.99 |  | 299.86 | Tare
| 299.99 |  | 299.88
| 299.99 |  | 299.88
| 299.99 |  | 299.88

## Sheet1
Electronic Components
Arduino Microcontroller (Nano) | 20.7 | 5.72
$20.70 [Arduino Store]
(Knock-off available with cable for $5.72 [Geekcreit])
USB-A to mini-B USB (or micro, or USB-B, depending on the specific Arduino in use) | 5.26 | 0
$5.26 [AmazonBasics]
5V USB Power Block (Optional) | 4
$4.00 [Amazon USB Wall Charger]
HX711 Loadcell Amplifier | 8.5
$8.50 with TAL220 [Amazon HX711]
Push Button (Normally Open, Momentary) | 2.5 | 0.125
$2.50 for 20 pack [Adafruit Tactile Button]
Jumper Wires
Breadboard:
Option 1: Solderless Breadboard (remains external to scale) | 5.95
$5.95 [Adafruit Breadboard]
Option 2: Solder Breadboard (40x60mm), Solder, Soldering Iron | 5.99 | 0.14975
$5.99 in 40-pack [Banggood]
LCD is optional. The three components below are required if using an LCD.
16x2 LCD Display | 5.99
$5.99 [Arducam]
10 kΩ Potentiometer | 4.38
$4.38 [DigiKey]
220 Ω Resistor | 7.95 | 0.0159
$7.95 in 500 pack [SparkFun]
If using a solderless breadboard, no more electronic components needed. Using a solder breadboard:
Female Header Pins: | 0.95
1x4 (2)
1x8 (1)
1x9 (1)
$0.95 [Adafruit Headers]
Male Header Pins: | 3 | 0.6
1x12 (1)
1x4 (1)
$3.00 for 5x16 [Adafruit Headers]
Configuration Specific Hardware
TAL 220 Loadcell configuration
TAL 220 Loadcell
Wires terminated in Female 4-pin molex connector with color order RKWG
$0.00 with HX711 [Amazon HX711]
Fasteners: | 2
M4x25mm cap screw (2) – 3mm Allen Key
M5x25mm cap screw (2) – 4mm Allen Key
<$2.00 from a hardware store
TAL 221 Loadcell configuration
TAL 221 Loadcell | 8.95
Wires terminated in Female 4-pin molex connector with color order RKWG
$8.95 [SparkFun]
Fasteners: | 2
M3x20mm cap screw (4) – 2.5mm Allen Key
M3 nut (2) – 5.5mm Socket
<$2.00 from a hardware store

## 200221 - 500g vs A-160 LC2
Mass Balance Comparison Test
| Device Under Test: |  | OS_Nano_Balance (3D printed balance), Using TAL221 500g load cell
| Lab Balance: |  | Denver A-160 160/0.1mg
| DUT Calibration: |  | A-160 measured plastic container to be 16.4059g
2889.79590 div/g
Added mass: instrument specialists part no 1011 hermettic lids |  |  |  |  | VWR Glass Vials with phenolic cap on Cat. No: 66011-041
n lids | APX402 (g) | A-160 (g) | DUT (g) | n vials | APX402 (g) | A-160 (g) | DUT (g) | Lab Balance (g) | DUT (g) | Lab Balance (g) | DUT (g)
6 | [5 empty] | 6.0118
Average | #DIV/0! | #DIV/0! |  |  | #DIV/0! | 6.0118 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Notes |  |  |  | Notes |  |  | Notes |  | Notes
Wiring is the issue - need to flip it around; mine is upside-down and the wires don't reach; will resolve these issues and re-test next week.

## 200221 - 500g vs A-160 LC1
Mass Balance Comparison Test
| Device Under Test: |  | OS_Nano_Balance (3D printed balance), Using TAL221 500g load cell
| Lab Balance: |  | Denver A-160 160/0.1mg
| DUT Calibration: |  | A-160 measured plastic container to be 16.4059g
2889.79590 div/g
Added mass: instrument specialists part no 1011 hermettic lids |  |  |  |  | VWR Glass Vials with phenolic cap on Cat. No: 66011-041
n lids | APX402 (g) | A-160 (g) | DUT (g) | n vials | APX402 (g) | A-160 (g) | DUT (g) | Lab Balance (g) | DUT (g) | Lab Balance (g) | DUT (g)
6 | 0.05 | 0.0801 |  | 1 | 6.01 | 6.0118 | 5.736
Average | 0.05 | 0.0801 |  |  | 6.01 | 6.0118 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Notes |  |  |  | Notes |  |  | Notes |  | Notes
| Value would not stay steady - cannot measure down to this precision |  |  |  | Just cannot get a good reading. Trying a new load cell

## 200121
Mass Balance Comparison Test
| Device Under Test: |  | OS_Nano_Balance (3D printed balance), Using TAL220 5000g load cell
| Lab Balance: |  | Digital Scale Waoaw 500x0.01g scale (from MOST Lab at MTU)
| DUT Calibration: |  | Measured the mass of a pocket knife using the Lab Balance, took the average mass, then used that as the calibration mass for the DUT.
| Pocket Knife |  |  | 3mm allen key (Lulzbot) |  |  | Partially empty filament roll
| Lab Balance (g) | DUT (g) |  | Lab Balance (g) | DUT (g) |  | Lab Balance (g) | DUT (g)
| 84.97 | 84.9 |  | 7.03 | 7 |  | 428.49 | 429.2 |  | 1002.3
| 85.1 | 84.9 |  | 7.02 | 7.1 |  | 428.49 | 429.3 |  | 1002.6 | 0.4
| 85.1 | 84.9 |  | 7.04 | 7 |  | 428.48 | 429.2 |  | 1002.8 | 0.7
| 84.88 | 85 |  | 7.03 | 7 |  | 428.46 | 429.4 |  | 1003.1 | 0.9
| 85.02 | 84.9 |  | 7.01 | 7 |  | 428.48 | 429.5 |  | 1001.8 | 1
| 84.96 | 84.8 | Tare.. | 7.02 | 6.9 | (balance restart for logging) | 428.48 | 429.1 | (tare…) | 1002.9 | 1.3
| 85.02 | 84.8 |  | 7.04 | 6.9 |  | 428.47 | 429.2 |  | 1003.3 | 1.4
| 85 | 84.7 |  | 7.01 | 7.1 | (tare) | 428.48 | 429.2 |  | 1001.8 | -0.1 | restart
| 84.99 | 84.8 |  | 7.03 | 7.1 |  | 428.49 | 429.3 |  | 1001.8 | -0.1
| 85.01 | 84.9 |  | 7.02 | 7.1 |  | 428.47 | 429.3 |  | 1001.8 | -0.1
Average | 85.005 | 84.86 |  | 7.025 | 7.02 |  | 428.479 | 429.27 |  | 1001.8
| Notes |  |  | Notes |  |  | Notes |  |  | Notes
| The mass does not always drop back to zero… seems to be drifting upward. I had to tare during the set of tests, which kinda skews the data. Highest I've seen the mass drift up is 0.2g.. Changing item placement on the balance has minimal effect (< 0.2g) on the measurement. | [5 empty] | Definitely need to be careful about contact with any wires. Ideally should clean up display wires to fall under the scale and out of the way. |  |  | Seem too be experiencing some variation in the returning point of the load cell… sometimes. Not quite hysteresis. It goes away for a while when restarting the system, then slowly starts building again. Slim chance that it has something to do with the acquisition software. Likely sensor temperature or something along those lines. |  |  | Sensor drift is temperature related. By blowing inside the balance the measured mass raised from zero to 40g, then drifted back down to zero (line 17928 of 20200121.log)